## A_Upfront_Considerati
A. UPFRONT CONSIDERATIONS (Germline)
Elements | Key Concepts |  | Additional Background and Educational Vignettes
TEST UTILITY (TYPE) | Diagnostic a) Disease-focused Gene Panel b) Exome/Genome Sequencing | Disease-focused Gene Panel • Typically used for "clear" clinical presentations (intended to detect a genetic cause for disease). Typically proband-only testing to confirm or rule out clinical diagnosis. • Ideally cover all genes that have well-supported disease relationships. If test size is limited, concessions can be made to exclude "minor" genes (genes with low detection rates). • Diagnostic testing typically requires higher degree of "completeness" with regard to the % of the gene for which data can be guaranteed and with regard to relevant variant types (eg, SNVs, indels, CNVs) that can be detected. For NGS-based tests, this typically translates to a higher average coverage to ensure that all bases are covered at or above the minimum coverage needed for reliable variant calling. • Disease focused gene panels usually require NGS technology (various approaches, typically capture- or amplicon- based). The assay may be limited to only the genes that are relevant for the disease of interest but laboratories increasingly design multi-disease assays or use an exome/genome panel as the base assay to maximize effiency for test development, operations, and revising content (the same assay can be used for many gene panels). However, exome-/genome- based panels typically achieve lower average coverage compared to disease-focused assays, which may affect the degree of completeness. • Variants included on clinical reports typically include variants classified as variant of uncertain significance (VUS), likely pathogenic, and pathogenic. Exome/Genome Sequencing • Typically used for clinically complex presentations of suspected genetic origin or "diagnostic odysseys." • Testing often includes additional family members for added clinical sensitivity (parent/child trio, multiple affected individuals). • Genes seleted "ad hoc" based on clinical features - may include candidate genes. • Degree of completeness usually not as stringent as for diagnostic gene panels. • Variants included on clinical reports are typically restricted to variants classified as likely pathogenic and pathogenic, sometimes VUSs that are relevant to the clinical presentation. • Secondary findings (ie, variants in "actionable" genes unrelated to the patient's clinical features) are an optional add-on to be reported. | Diagnostic Gene Panels: • PMID: 31732716 (Diagnostic Gene Sequencing Panels: From Design to Report-A Technical Standard of the American College of Medical Genetics and Genomics (ACMG) Diagnostic Whole Exome/Genome sequencing: • PMID: 28362156 (Development and validation of clinical whole-exome and whole-genome sequencing for detection of germline variants in inherited disease) Carrier Screening: • PMID: 28225426 (ACOG Committee Opinion No. 691: Carrier Screening for Genetic Conditions) • PMID: 28225425 (Committee Opinion No. 690: Carrier Screening in the Age of Genomic Medicine) Elective Genomic Testing: • PMID 30453057 (Evaluation for Genetic Disorders in the Absence of a Clinical Indication for Testing). Elective Genomic Testing: for an educational review of considerations around analytical and interpretative components of genomic tests. PLEASE NOTE: • Additional considerations around variant interpretation and reporting for several of the above test types is provided in a separate worksheet (interpretation and reporting) | Exome Slice Testing - Considerations from Ordering to Reporting: A Report of the Association for Molecular Pathology
| Carrier Screening | A genetic test that is used to determine if a person is a carrier for specific autosomal recessive diseases. This kind of testing is used most often by couples who are considering becoming pregnant to determine the risks of their child inheriting one of these genetic disorders.
| Elective Genomic Testing | An emerging type of test that is offered to healthy individuals who do not have a clinical indication. Because these tests can have significant effect on medical decision making, they typically require higher stringency for selecting genes with an increased focus on the validity of gene-disease relationship compared to diagnostic applications and as well as a greater focus on actionability. Additionally, reporting is typicaly limited to pathogenic and likely pathogenic variants.
TEST INDICATION (USE) | Use case | • Which disorder/disease/condition is this test intended for. • Which specific population will this test be offered to (eg, which age range, etc) | Example: • Test indication: hypertrophic cardiomyopathy (HCM) • Specific population: test typically offered for all ages
SPECIMEN TYPES | Sample types | • Determine all specimen types that will need to be accepted for optimal utility. • Validation must accommodate each specimen type. | Post-mortem diagnostic testing for sudden death: in this setting, the predominant specimen type currently is FFPE tissue, which tends to perform at a lower level compared to whole blood and saliva specimens. Newborn testing: commonly performed using dried blood spots. Diagnostic and carrier screening: most common sample type used is peripheral blood
TEST CONTENT | Genes | • What genes, and regions, the test is supposed to cover based on the test indication. • Are there overlapping clinical diagnoses that must be covered for maximum clinical utility (ie, what is the differential diagnosis, the list of conditions that can manifest with similar features) • The differential diagnosis can inform the list of genes that should be interpreted. | Diagnostic test for hypertrophic cardiomyopathy (HCM): • Include genes associated with HCM, isloated left ventricular hypertrophy (LVH), and genes associated with syndromic conditions that can present with apparently isolated LVH and can therefore be misdiagnosed when LVH is the first or only obvious feature (general guidance: PMID: 31732716) • Expert-curated gene-disease validity levels for HCM (Ingles 2019, PMID: 30681346)
| Variant types | • Which variant types are clinically relevant (ie, are known to contribute to disease)? • What are their relative contributions to disease? • Will all types need to be detected at the same level of analytical sensitivity? | • For optimal clinical validity, the test should detect all variant types that are represented among disease causing variants. • For inherited disorders this usually includes SNVs and indels at a minimum. • When CNVs are an important variant type, additional detection and/or confirmation platforms may need to be added to the development plan.
TECHNOLOGY LIMITATIONS | Interrogated regions | At a high level, are there well documented technical issues with regard to the planned content? This can often be accomplished by a literature review: • Such information is often readily available in review articles for genetically well-defined disorders for which testing has been offered for a long time. • Technical issues include high homology to pseudogenes, high or low GC content, the existence of difficult to detect variants types such as repeat expansions, etc. • The existence of such issues informs the likelihood for requiring supplemental assays to ensure maximum analytical sensitivity. For example, if a critical gene cannot be sequenced by short fragment NGS due to high sequence homology, another assay needs to be developed and included in the overall test. | Diagnostic testing for inherited sensorineural hearing loss (SNHL): • A gene panel should include the stereocilin (STRC) gene, which is responsible for ~10% of SNHL. This gene has high sequence homology to a pseudogene and the pathogenic variant spectrum includes small variants (SNVs and indels) as well as large multi-exon to whole gene deletions). Alternate/supplemental assays that can complement NGS need to be considered (PMID: 25157971). Diagnostic testing for inherited kidney disease: • Pathogenic variants in the PKD1 gene are responsible for 85% of autosomal dominant polycystic kidney disease (AD-PKD). However, this gene is highly homologous to several pseudogenes, making analysis via standard short read NGS difficult. In this case, a stand alone Sanger sequencing test can be used as a first tier test. The clinical specificity of AD-PKD is high such that a broad gene panel is less crucial and can be used as a reflex option for negative cases. General guidance on how to navigate test design and development for genes with technical complexity: • Santani et al. 2017 (PMID: 28322587)
INTERFERING SUBSTANCES |  | • There are two general sources of interfering substances, exogenous and endogenous. Exogenous interfering substances are introduced to the specimen through various routes (eg, pharmacological therapy, substances present in specimen collection tubes, food present in saliva specimens). Endogenous interfering substances can be normal components of metabolism or the result of pathological conditions (eg, lipids, bilirubin, hemoglobin, or albumin). • For germline testing, interfering substances are less of a concern when whole blood is the specimen type. However, the use of saliva as a specimen type is increasing and exogenous contamination by either microbial content and/or food remnants introduces interference that must be considered upfront. | Food and microbial content present in saliva specimens: • Non-human DNA contamination will not be detected during DNA extraction and can lead to a test failure due to insufficient alignable reads and/or a high fraction of low-quality variant calls. • Additional analytical steps (eg, a "pre-run" + alignment against bacterial genomes) may be needed to avoid high failure rates and effects on TATs.
ANALYTICAL PERFORMANCE NEEDS | Read depth | • This is an important consideration for tests that require a high degree of completeness (see test type, above). • Does the test developer plan on filling in missing data if required completeness cannot be achieved by the assay (eg, by using Sanger sequencing to "fill in" regions that have coverage below the required minimum coverge for reliable variant detection?). | If a high degree of completeness is essential, the laboratory needs to plan on whether and how to use fill-in technologies (such as Sanger sequencing): • This is influenced by the specific capabilities and existing testing, which varies greatly between laboratories. If the fill-in technology is not yet established in the laboraotry, the cost of doing so must be taken into account. Generally speaking, the laboratory should anticipate a need to achieve higher average coverage to ensure that all/most bases are covered at or above the needed minimum coverage. This affects the number of samples/sequencing run and therefore cost.
| Analytical sensitivity (Positive percent agreement) (TP/[TP+FN]) | • Test ability to detect all confirmed variants. • Is there a minimum acceptable analytical sensitivity that must be achieved? • The test developer should predefine target thresholds including the lower bound of the 95% confidence interval. • Test optimization will need to be iterated until this is achieved for all clinically relevant variant types. | • FDA: Considerations for design, development, and analytical validation of Next Generation Sequencing (NGS)-based in Vitro Diagnostics (IVDs) intended to aid in the diagnosis of suspected germline diseases: (https://www.fda.gov/downloads/MedicalDevices/DeviceRegulationandGuidance/GuidanceDocuments/UCM509838.pdf) • Generally, optimal clinical validity requires near-perfect analytical sensitivity. This is impacted by various factors such as including the type of variant caller used and read depth. • Applications that may tolerate a reduced analytical sensitivity are emerging testing areas where no other test options exist. • See see applicable parts of 21 CFR part 801 and 21 CFR 809.10 (https://www.ecfr.gov/cgi-bin/text-idx?SID=3ee286332416f26a91d9e6d786a604ab&mc=true&tpl=/ecfrbrowse/Title21/21tab_02.tpl).
| Analytical specificity (Negative percent agreement) (TN/[TN+FP]) | • Test ability to correctly identified WT bases in the interrogate region. • Does the test developer plan on confirming variants using a 2nd assay? | • YES: A lower NGS specificity is acceptable but the laboratory must take additional cost incurred into account. • NO: NGS specificity must be high. Limitations must clearly state that additional confirmation should be sought prior to using the results for medical action.
| Technical Positive Predictive Value (PPV) (TP/[TP+FP]) | • Likelihood that an NGS variant call is a true positive. • The test developer should predefine thresholds for variant call metrics to achieve the best technical PPV for a defined region of interes, without affecting the analytical sensitivity. • It is important to calculate the PPV for each variant types because you cannot accurately determine the specificy of an NGS assay for all variants discovery. | Example: • The test assesses 100,000 base pairs, 200 variants are identified of which 5 are FP • Specificity = 99,995/100,000 = 99.99% • PPV = 200/205 = 97.56%
CLINICAL PERFORMANCE NEEDS |  | • Is the test well-established in the community and has clinical validity (genotype and clinical phenotype correlation). • Do performance expectations exist or can a lower clinical sensitivity (detection rate, the % of cases that will receive a positive result) or clinical specificity (the ability of the test to correctly identify those patients without the disease) be acceptable?
FEASIBILITY ASSESSMENT | Sample volume | • Is there a predetermined number of samples the test developer must be able to process in a given time frame? | Example: • The test is being developed for the cardiology department of a specific hospital. • This department has an existing patient volume of 500/month and expects to order 20 gene panels for inherited cardiomyopathies every month. Key areas impacted (typically taking turnaround time requirements into consideration): • Type of sequencer • Testing personnel projections (number of FTEs to accommodate volume)
| Turnaround time | • Are there clinical considerations that inform the turnaround time (TAT)? • If yes, does the laboratory expect to receive the necessary sample volume to reliably return results within this TAT while meeting cost targets? | A need for a fast TAT is typically encountered with: • Prenatal testing for hereditary disorders • Hereditary cancer testing to tailor surgery options • Diagnostic testing for a fast progressing disease with defined treatment path
| Cost | • Is the test reimbursed? At what level? • Does the laboratory handle 3rd party payment? If yes, will the reimbursement level meet the laboratory's financial management goals?
REGULATORY | Regulatory agencies | • Regulatory requirement from certain agencies such as the FDA in the USA and ISO in Europe that have different analytical requirements should be considered
IS IT A GO? | YES or NO
© 2024 College of American Pathologists (CAP). All rights reserved.
CLSI. Human Genetic and Genomic Testing Using Traditional and High-Throughput Nucleic Acid Sequencing Methods. 3rd ed. CLSI guideline MM09. Clinical and Laboratory Standards Institute; 2023.

## References
Year | Entity | PMID | Title
General Recommendations and Guidelines
2019 | CCMG | 31300550 | CCMG practice guidance: laboratory guidelines for next-generation sequencing.
2018 | FDA | N/A | Use of Public Human Genetic Variant Databases to Support Clinical Validity for Genetic and Genomic-Based In Vitro Diagnostics https://www.fda.gov/regulatory-information/search-fda-guidance-documents/use-public-human-genetic-variant-databases-support-clinical-validity-genetic-and-genomic-based-vitro.
2017 | CDC | 28315672 | Principles and Recommendations for Standardizing the Use of the Next-Generation Sequencing Variant File in Clinical Settings.
2017 | EU | N/A | REGULATION (EU) 2017/746 OF THE EUROPEAN PARLIAMENT AND OF THE COUNCIL of 5 April 2017 on in vitro diagnostic medical devices and repealing Directive 98/79/EC and Commission Decision 2010/227/EU (https://eur-lex.europa.eu/legal-content/EN/TXT/PDF/?uri=CELEX:32017R0746&rid=6.)
2016 | ESHG EuroGentest | 26508566 | Guidelines for diagnostic next-generation sequencing.
2015 | CAP | 25152313 | College of American Pathologists' laboratory standards for next-generation sequencing clinical tests.
2014 | CAP | 24650895 | Methods-Based Proficiency Testing in Molecular Genetic Pathology.
2012 | CDC | 23138292 | Assuring the quality of next-generation sequencing in clinical laboratory practice.
2010 | EuroGentest | 20664632 | A standardized framework for the validation and Verification of Clinical Molecular Genetic Tests.
N/A | New York State | N/A | New York State Department of Health, Wadsworth Center. Test Approval. https://www.wadsworth.org/regulatory/clep/clinical-labs/obtain-permit/test-approval.
Inherited Disorders
2023 | AMP | 38103592 | Slice Testing-Considerations from Ordering to Reporting: A Joint Report of the Association for Molecular Pathology, College of American Pathologists, and National Society of Genetic Counselors
2020 | ACMG | 31732716 | Diagnostic Gene Sequencing Panels: From Design to Report - A Technical Standard of the American College of Medical Genetics and Genomics (ACMG)
2019 | CAP | 30605766 | Designing and Implementing NGS Tests for Inherited Disorders: A Practical Framework with Step-by-Step Guidance for Clinical Laboratories.
2018 | FDA | N/A | Considerations for design, development, and analytical validation of Next Generation Sequencing (NGS)-based in Vitro Diagnostics (IVDs) intended to aid in the diagnosis of suspected germline diseases: (https://www.fda.gov/downloads/MedicalDevices/DeviceRegulationandGuidance/GuidanceDocuments/UCM509838.pdf)
2017 | CAP | 28322587 | Development and validation of targeted next-generation sequencing panels for detection of germline variants in inherited diseases.
2017 | CAP | 28362156 | Development and validation of clinical whole-exome and whole-genome sequencing for detection of germline variants in inherited disease.
2020 | New York State | N/A | Next-Generation Sequencing (NGS) guidelines for germline genetic variant detection . https://www.wadsworth.org/sites/default/files/WebDoc/Germline_NextGen_Validation_Guidelines_FINAL_2020_rev_09092021.pdf
2013 | ACMG | 23887774 | ACMG clinical laboratory standards for next-generation sequencing.
2012 | AMP | 22918138 | Opportunities and challenges associated with clinical diagnostic genome sequencing: A report of the Association for Molecular Pathology.
Oncology
2021 | New York State | N/A | Next Generation Sequencing (NGS) Guidelines for somatic genetic variant detection. https://www.wadsworth.org/sites/default/files/WebDoc/NextGenSeqONCOGuidelines%20_April_2021.pdf
2017 | ACMG/AMP/CAP | 27993330 | Standards and Guidelines for the Interpretation and Reporting of Sequence Variants in Cancer: A Joint Consensus Recommendation of the Association for Molecular Pathology, American Society of Clinical Oncology, and College of American Pathologists.
2017 | AMP/CAP | 28341590 | Guidelines for Validation of Next-Generation Sequencing-Based Oncology Panels: A Joint Consensus Recommendation of the Association for Molecular Pathology and College of American Pathologists.
2012 | CLSI | N/A | MM01: Molecular Methods for Clinical Genetics and Oncology Testing, 3rd Edition
Bioinformatics
2018 | AMP/CAP | 29154853 | Standards and guidelines for validating next-generation sequencing bioinformatics pipelines: a joint recommendation of the Association for Molecular Pathology and the College of American Pathologists.
2015 | CDC | 26154004 | Good laboratory practice for clinical next-generation sequencing informatics pipelines.