## BEFORE YOU BEGIN_V3.0
Before you begin: How to use this workbook
1. Start on Tab B, Gene_Disorder_Variant_Info
This sheet is intended for both manufacturers of in-vitro diagnostic (IVD) products and laboratory-developed test (LDT) developers and is designed to detail the genes, diseases, and critical variants you are trying to capture in your assay. It indicates which disorders you will report when there are multiple disorders associated with a gene, the mutational spectrum for variants in this gene associated with that disorder, and the prevalence of that gene in association with a particular disorder. For determining the strength of evidence for a gene-disorder association, following standard guidelines, such as those put forth by ClinGen (https://www.clinicalgenome.org/working-groups/gene-curation/projects-initiatives/clinical-validity-classifications/). Please note, although gene validity and mutation spectrum may be assessed at the interpretation and reporting stages for exome testing, the laboratory should have a general policy to consider these aspects.
2. Continue to Tab C, Target_Region_Technical_Info
This sheet is designed to capture specific information around each gene in terms of important technical considerations when designing the capture methodology or target content. Specifically, the aim is to predetermine regions for critical genes that may cause issues during the design, validation, or operation of a clinical NGS assay. Regions that commonly cause technical issues include regions of high homology or high or low GC content. The laboratory should establish thresholds for high or low GC content above or below which test performance may be affected. Difficult regions (such as highly homologous regions) can be determined through a variety of mechanisms. Examples include: a) Mandelker et al (https://www.ncbi.nlm.nih.gov/pubmed/27228465) b) GA4GH Benchmarking Tool team (https://github.com/ga4gh/benchmarking-tools/tree/master/resources/stratification-bed-files)
3. Continue to Tab D, Problem_List
This tab is intended to be used to integrate information assembled in tabs B and C. For example, information on tab C will enable flagging of regions with anticipated technical challenges. Cross referencing this information with the information assembled on Tab B (Gene_Disorder_Variant_Info) enables identification of problematic regions that are critical for clinical validity (eg, due to a high gene contribution or the existence of common pathogenic variants in a certain region). This analysis is performed at the exon level and enables the laboratory to determine whether problematic regions can be dropped from the test content or whether ancillary assays are required.
4. Continue to Tab E, Reference_Material_Selection
Once the test has been designed, the developer can use the information in Tab B (Gene_Disorder_Variant_Info) to assemble reference materials for assay optimization and subsequent test validation.
5. Tab F provides a list of reference material resources
© 2024 College of American Pathologists (CAP). All rights reserved.
CLSI. Human Genetic and Genomic Testing Using Traditional and High-Throughput Nucleic Acid Sequencing Methods. 3rd ed. CLSI guideline MM09. Clinical and Laboratory Standards Institute; 2023.

## B_ Gene_Disorder_Variant_Info_V
B. Gene-Disorder Variant Information
SECTION 1
This Tab is designed to help: | [9 empty] | Resource | URL (click to access directly)
• Assemble candidate gene list for clinical indication served by the test | [9 empty] | • Relevant Professional Standards | ACMG Standard for Dx Panel design (2020)
• Determine the strengh of gene-disease association to ensure optimal clinical validity | [9 empty] | • Test Indication and Use | See Test Familiarization Worksheet
• Determine the maximum possible detection rate across clinically valid genes (clinical sensitivity) | [9 empty] | • HGNC Gene symbol Lookup | HUGO Gene Nomenclature Lookup
• Determine the spectrum of clinically relevant variant types to inform assay and pipeline needs + ensure optimal analytical validity | [9 empty] | • ClinGen Gene-Disease Validity | ClinGen Website - LINK
• Assemble a list of prevalent pathogenic variants to inform validation needs and reference material selection | [9 empty] | • GenCC Gene-Disease Validity | GenCC
• Define the clinical "Region of Interest" (ROI), defined as the sum of regions that need to be part of the test for optimal clinical validity | [9 empty] | • MANE Select transcript set | MANE Select Info Page (NCBI)
SECTION 2 (SUMMARY OF SECTION 3 CONTENT)
ROI Default Settings (use along with specifications in Table 1 to build Target Region and Annotate in Tab C) | [9 empty] | REPRESENTATIVE ROI (REGION OF INTEREST) DEFAULTS
Coding exons | See educational example in Figure to the right
Patially coding exons
Non co-coding exons
Other functional intervals (eg promoter)
Note: The concept of 'region of interest" is mentioned throughout MM09. The details listed here are not included in the current published version.
SUMMARY |  |  | Data pulled in from Table 1 and Table 2 below
Which Clinical entities need to be included? (clinical overlap; differential diagnosis) | Disease Code | Prevalence | Expected clinical sensitivity (Table 2) | Clinically Significant Variant Spectrum (Use to inform assay/pipeline design ) |  |  |  | Recurrent Prevalent Variants (Y/N) Add variants to RM selection (Tab E)
|  |  |  | SNV + InDel | CNV | Other SV | Other (eg: rpt exp)
disease 1 | OMIM: 123456 | 1/500 | 56.8% | Y | Y | N | N
disease 2 | OMIM: 345678 | 1/300 | 48.66% | Y | N | N | N
disease 3 | add | add | formula --> table 2 | --> table 1 | --> table 1 | --> table 1 | --> table 1
disease 4 | add | add | formula --> table 2 | --> table 1 | --> table 1 | --> table 1 | --> table 1
disease 5 | add | add | formula --> table 2 | --> table 1 | --> table 1 | --> table 1 | --> table 1
disease 6 | add | add | formula --> table 2 | --> table 1 | --> table 1 | --> table 1 | --> table 1
Courtesy of: Annette Leon and Birgit Funke
SECTION 3 (DETAILED ANNOTATIONS)
TABLE 1: GENE - DISEASE DATA | [5 empty] | TABLE 2: DETECTION RATES
Gene | GENE 1 | GENE 2 | GENE 3 | GENE N |  | GENE | DISEASE 1 |  | DISEASE 2
HGNC Symbol (See HGNC LookUp URL above) | [6 empty] | Exp. Det. Rate | Sources | Exp. Det. Rate | Sources
Alias (eg legacy name) | [6 empty] | 56.8% |  | 48.66%
Ensembl ID | [5 empty] | GENE01 | 1% |  | 0%
Disease | DISEASE 1 | DISEASE 2 | DISEASE 3 | DISEASE N |  | GENE02 | 4% |  | 0%
Strength of Gene-Disease Association (ClinGen if available; GenCC is a good alternative) URLs: See above | DEFINITIVE | DEFINITIVE | DEFINITIVE |  |  | GENE03 | 2% |  | 0%
[6 empty] | GENE04 | 2% |  | 0%
Inheritance Pattern | AD | AD | AD |  |  | GENE05 | 1.5% |  | 0%
Molecular Mechanism | LOF | GOF | GOF |  |  | GENE06 | 10% |  | 0%
Relative Contribution of Variant Types --> use to fill summary table above + inform assay/pipeline design | [5 empty] | GENE07 | 4.5% |  | 0%
SNV | 5% | 99% | 1% |  |  | GENE08 | 5% |  | 0%
InDel | 80% | 1% | 99% |  |  | GENE09 | 1% |  | 0%
CNV | 15% | 0% | 0% |  |  | GENE10 | 0.1% |  | 0%
Other SV | 0% | 0% | 0% |  |  | GENE11 | 2% |  | 0%
Other (eg repeat expansion) | 0% | 0% |  |  |  | GENE12 | 0.3% |  | 0%
Recurrent/Prevalent Variants --> --> --> add to Tab E (reference material selection) | [5 empty] | GENE13 | 1% |  | 0%
coding | [5 empty] | GENE14 | 1% |  | 0%
Variant 1 | [5 empty] | GENE15 | 0% |  | 0%
Variant 2 | [5 empty] | GENE16 | 15% |  | 0%
Variant N | [5 empty] | GENE17 | 0.4% |  | 0%
Etc | [5 empty] | GENE18 | 1% |  | 3%
non-coding | [5 empty] | GENE19 | 4.9% |  | 10.56%
Variant 1 | [5 empty] | GENE20 | 0.1% |  | 5%
Variant 2 | [5 empty] | GENE21 | 0% |  | 3%
Variant N | [5 empty] | GENE22 | 0% |  | 10%
Etc | [5 empty] | GENE23 | 0% |  | 2%
[6 empty] | GENE24 | 0% |  | 5%
Clinically Relevant Transcripts (known relevance for disease - OR - all transcripts if nothing is known) | [5 empty] | GENE25 | 0% |  | 1%
Primary transcript for reporting (choose one option) | [5 empty] | GENE26 | 0% |  | 1%
| Transcript ID |  |  |  |  | GENE27 | 0% |  | 1%
• MANE Select (if available) (URL above) | add ID |  |  |  |  | GENE28 | 0% |  | 5%
• Laboratory defined | add ID |  |  |  |  | GENE29 | 0% |  | 2%
[6 empty] | GENE30 | 0% |  | 0.1%
Additional transcripts (alternatively spliced exons)
Source | Transcript ID | alt exon(s) | coding?
• MANE Plus Clinical (if available) |  | ex 1, 6, 7 | N/Y/Y
• Laboratory defined |  | ex 5, 20 | Y/Y
Other Regions of known clinical relevance
Region | coordinates
e.g.: promoter | add
e.g.: intronic splice branch site | add

## C_Target_Region_Technical_Info_
C. Target Region Technical Information
1. Gene Regions | [10 empty] | 2. Technical Considerations |  | 3. Test Region of Interest |  |  |  |  | 4. Region Flags
Gene ID | Transcript | Genome Build | Genomic Coordinates | Region Size (bp) | Strand | Region Name | Reason for Inclusion | Reason for Custom Padding | cDNA Start | cDNA End | % GC | % High Homology (low MQ bases) | Chr | Start | End | Strand | Gene_Region | Custom Padding? (Yes/No) | GC Content Issue? (Yes/No) | Homology Issue? (Yes/No) | Critical Variants? (Yes/No)
BSND | NM_057176.2 | GRCh37 | chr1:55464855-55465041 | 187 | + | Exon 01 | coding exon | default | c.-5 | c.177+5 | 63.25% | 0% | 1 | 55464855 | 55465041 | + | BSND_Exon_01 | No | No | No | No
BSND | NM_057176.2 | GRCh37 | chr1:55470690-55470794 | 105 | + | Exon 02 | coding exon | intronic variant | c.178-25 | c.272+5 | 59.63% | 0% | 1 | 55470690 | 55470794 | + | BSND_Exon_02 | Yes | No | No | No
BSND | NM_057176.2 | GRCh37 | chr1:55472665-55472950 | 286 | + | Exon 03 | coding exon | default | c.273-5 | c.548+5 | 59.87% | 0% | 1 | 55472665 | 55472950 | + | BSND_Exon_03 | No | No | No | No
BSND | NM_057176.2 | GRCh37 | chr1:55473882-55474306 | 425 | + | Exon 04 | coding exon | default | c.549-5 | c.*5 | 58.91% | 0% | 1 | 55473882 | 55474306 | + | BSND_Exon_04 | No | No | No | No
EDNRB | NM_001201397.1 | GRCh37 | chr13:78493527-78493755 | 229 | - | Exon 01 | coding exon | default | c.-5 | c.219+5 | 56.18% | 0% | 13 | 78493527 | 78493755 | - | EDNRB_Exon_01 | No | No | No | No
EDNRB | NM_001201397.1 | GRCh37 | chr13:78492221-78492764 | 544 | - | Exon 02 | coding exon | default | c.220-5 | c.753+5 | 57.95% | 0% | 13 | 78492221 | 78492764 | - | EDNRB_Exon_02 | No | No | No | No
EDNRB | NM_001201397.1 | GRCh37 | chr13:78477625-78477747 | 123 | - | Exon 03 | coding exon | default | c.754-5 | c.866+5 | 42.14% | 0% | 13 | 78477625 | 78477747 | - | EDNRB_Exon_03 | No | Yes | No | No
EDNRB | NM_001201397.1 | GRCh37 | chr13:78477286-78477500 | 215 | - | Exon 04 | coding exon | default | c.867-5 | c.1071+5 | 36.4% | 0% | 13 | 78477286 | 78477500 | - | EDNRB_Exon_04 | No | Yes | No | No
EDNRB | NM_001201397.1 | GRCh37 | chr13:78475188-78475347 | 160 | - | Exon 05 | coding exon | default | c.1072-5 | c.1221+5 | 30.22% | 0% | 13 | 78475188 | 78475347 | - | EDNRB_Exon_05 | No | Yes | No | No
EDNRB | NM_001201397.1 | GRCh37 | chr13:78474651-78474794 | 144 | - | Exon 06 | coding exon | default | c.1222-5 | c.1355+5 | 40.76% | 0% | 13 | 78474651 | 78474794 | - | EDNRB_Exon_06 | No | Yes | No | No
EDNRB | NM_001201397.1 | GRCh37 | chr13:78473989-78474107 | 119 | - | Exon 07 | coding exon | default | c.1356-5 | c.1464+5 | 33.19% | 0% | 13 | 78473989 | 78474107 | - | EDNRB_Exon_07 | No | Yes | No | No
EDNRB | NM_003991.3 | GRCh37 | chr13:78472330-78472474 | 145 | - | Exon 07B | coding exon | default | c.1465-5 | c.*5 | 35.8% | 0% | 13 | 78472330 | 78472474 | - | EDNRB_Exon_08 | No | Yes | No | No
EDNRB | NM_001201397.1 | GRCh37 | chr13:78470572-78470698 | 127 | - | Exon 08 | coding exon | default | c.1195-5 | c.*5 | 35.8% | 0% | 13 | 78470572 | 78470698 | - | EDNRB_Exon_07B | No | Yes | No | No

## D_Problem_List_V3.0
D. Problem List
1. Gene-Level Summary
Gene | # Exons GC Content Issues (<40% or >70%) | # Exons w/ Homology Issues | Total # Exons Affected | Total # Exons in Gene | % Gene Affected | Specific Difficult Variant(s) | Gene Essential for Clinical Validity/Utlity (Yes/No) | If Yes, Supplemental Assay/Method Needed?
STRC | 0 | 20 | 20 | 25 | 80% | No | Yes | Long Range PCR
PHOX2B | 0 | 0 | 0 | 5 | 0% | Poly Ala repeat (EX03) | Yes | Sanger Sequencing
EDNRB | 7 | 0 | 7 | 9 | 77.78% | No | Yes | No
2. Exon-Level Summary
Gene | Transcript | Region Size (bp) | Exon/Region Name | Reason for Inclusion | GC Content Issue? (Yes/No) | Homology Issue? (Yes/No) | Critical Variants? (Yes/No) | Include in Assay? (Yes/No)
EDNRB | NM_001201397.1 | 123 | Exon 03 | coding exon | Yes | No | No | Yes
EDNRB | NM_001201397.1 | 215 | Exon 04 | coding exon | Yes | No | No | Yes
EDNRB | NM_001201397.1 | 160 | Exon 05 | coding exon | Yes | No | No | Yes
EDNRB | NM_001201397.1 | 144 | Exon 06 | coding exon | Yes | No | No | Yes
EDNRB | NM_001201397.1 | 119 | Exon 07 | coding exon | Yes | No | No | Yes
EDNRB | NM_003991.3 | 145 | Exon 07B | coding exon | Yes | No | No | Yes
EDNRB | NM_001201397.1 | 127 | Exon 08 | coding exon | Yes | No | No | Yes

## E_Reference_Material_Selection_
E. Reference Material Selection
Checklist for Assembling Reference Materials:
1. Assemble reference samples according to requirements established during test design (see Assay Design Optimization worksheet)
2. Determine numbers of variants by type
3. Are numbers sufficient for adequate statistical confidence (validation)?
4. All essential disease/gene specific requirements met?
5. Do samples need to be supplemented with synthetic materials to boost numbers?
Confidence Interval Calculator | http://www.pedro.org.au/english/downloads/confidence-interval-calculator/.
1. Summary of variants present in reference samples
Clinically Relevant Variant Spectrum | Variant Type Prevalent? See: Tab B: Gene_Disease_Variant_Info (Yes/No) | Total # in Reference Materials List (see table below) | Max possible lower bound of analytical sensitivity (95% CI) | Additional Detail
SNV | Yes | 606 | 0.9937
Indel (1-n base pairs) | Yes | 57 | 0.9369
CNV | Yes | 43 | 0.918 | Intragenic CNVs common
Other SV | No | 0 | N/A
Specific pathogenic variants | Yes | 13 | N/A | Variant 1, variant 2, variant 3, etc
Specimen Type | Required? (Yes/No) |  | Other Requirements | Required? (Yes/No)
Saliva | Yes |  | Somatic mosaicism | Yes
Whole blood | Yes |  | Heteroplasmy (mitoch) | No
Amniotic fluid | No |  | Other | No
CVS | No
DNA (prev. tested internal specimen) | Yes
DNA (extracted outside of testing lab) | Yes
2. Number of Variants in Test Region Per Reference Sample
Reference Sample | Specimen Type | Source | Validation Category a) Methods-based b) Gene/disease specific | Validation Purpose* | SNV | Indel (1-x bp) | Indel (x - y bp) | Indel (y - z bp) | CNV - DEL (<5 exons) | CNV - DEL (>= 5 exons) | Other SV | List Specific Pathogenic Variant/s
NA12878 | Extracted DNA | NIST | Methods-based | universal reference sample | 600 | 40 | 10 | 2 | 2 | 0 | 0
Sample 1 | Extracted DNA | Prev. tested specimen | Methods-based | CNV |  |  |  |  | 1 | 0
Sample 2 | Extracted DNA | Prev. tested specimen | Methods-based | CNV |  |  |  |  | 3 | 0
Sample 3 | Extracted DNA | Prev. tested specimen | Methods-based | CNV |  |  |  |  | 10 | 0
Sample 4 | Extracted DNA | Prev. tested specimen | Methods-based | CNV |  |  |  |  | 1 | 0
Sample 5 | Extracted DNA | Prev. tested specimen | Methods-based | CNV |  |  |  |  | 3 | 1
Sample 6 | Extracted DNA | Prev. tested specimen | Methods-based | CNV |  |  |  |  | 5 | 1
Sample 7 | Extracted DNA | Prev. tested specimen | Methods-based | CNV |  |  |  |  | 1 | 1
Sample 8 | Extracted DNA | Get-RM | Methods-based | CNV | [5 empty] | 1
Sample 9 | Extracted DNA | Get-RM | Methods-based | CNV | [5 empty] | 1
Sample 10 | Extracted DNA | Get-RM | Methods-based | CNV | [5 empty] | 1
Sample 11 | Extracted DNA | Get-RM | Methods-based | CNV | [5 empty] | 1
Sample 12 | Extracted DNA | Get-RM | Methods-based | CNV | [5 empty] | 1
Sample 13 | Extracted DNA | Get-RM | Methods-based | CNV | [5 empty] | 1
Sample 14 | Extracted DNA | Get-RM | Methods-based | CNV | [5 empty] | 1
Sample 15 | Extracted DNA | Get-RM | Methods-based | CNV | [5 empty] | 1
Sample 16 | Extracted DNA | Get-RM | Methods-based | CNV | [5 empty] | 1
Sample 17 | Extracted DNA | Get-RM | Methods-based | CNV | [5 empty] | 1
Sample 18 | Extracted DNA | Get-RM | Methods-based | CNV | [5 empty] | 1
Sample 19 | Extracted DNA | Get-RM | Methods-based | CNV | [5 empty] | 1
Sample 20 | Extracted DNA | Get-RM | Gene/disease specific | Prevalent pathogenic variants | 1 | [6 empty] | Variant 1
Sample 21 | Extracted DNA | Get-RM | Gene/disease specific | Prevalent pathogenic variants |  | 1 | [5 empty] | Variant 2
Sample 22 | Extracted DNA | Get-RM | Gene/disease specific | Prevalent pathogenic variants |  |  | 1 |  |  |  |  | Variant 3
Sample 23 | Extracted DNA | Coriell | Gene/disease specific | Prevalent pathogenic variants |  |  |  |  | 1 |  |  | Variant 4
Sample 24 | Extracted DNA | Coriell | Gene/disease specific | Prevalent pathogenic variants | [5 empty] | 1 |  | Variant 5
Sample 27 | Pooled synthetic ctrl | SeraCare | Gene/disease specific | Prevalent pathogenic variants | 5 | 1 | 1 | 1 |  |  |  | Variant 6-13
Sample 28 | Extracted DNA | Prev. tested specimen | Gene/disease specific | Somatic Mosaicism
Sample 29 | Extracted DNA | Prev. tested specimen | Gene/disease specific | Somatic Mosaicism
Sample 30 | Extracted DNA | Prev. tested specimen | Gene/disease specific | Somatic Mosaicism
Sample 31 | Extracted DNA | Prev. tested specimen | Gene/disease specific | Somatic Mosaicism
Sample 32 | Extracted DNA | Prev. tested specimen | Gene/disease specific | Somatic Mosaicism
Sample 32 | Extracted DNA (mix) | 10% variant/90% WT | Gene/disease specific | Somatic Mosaicism
Sample 33 | Extracted DNA (mix) | 20% variant/90% WT | Gene/disease specific | Somatic Mosaicism
Sample 34 | Extracted DNA (mix) | 30% variant/90% WT | Gene/disease specific | Somatic Mosaicism

## F_Reference_Material_Resources_
F. Reference Material Resources | Links
National Institute of Standards and Technology (commonly used samples are NA12878, NA24385, NA24385/NA24149/NA24143(trio), NA24631) | Genome in a Bottle ¦ NIST
NIST Standard Reference Materials ¦ NIGMS Collections ¦ Coriell
Genetic Testing Reference Materials Coordination Program (GeT-RM) | GeT-RM ¦ CDC
Coriell Institute for Medical Research | https://www.coriell.org/0/sections/collections/HuRef/Samples.aspx?PgId=644&coll=HR
European Quality Monitoring Network (EMQN) https://www.emqn.org/ | EMQN
The Human Reference Genetic Material Repository DNA Samples (Coriell) | The Human Reference Genetic Material Repository - DNA Samples
Multiplexed inherited cancer reference material (20 variants) | Seraseq Inherited Cancer DNA Mix
Multiplexed inherited cardiomyopathy reference material (10 variants) | Seraseq Inherited Cardiomyopathy DNA Mix
Multiplexed cystic fibrosis reference material (47 variants) | Seracare CF 40+4 RM (#2040-0015)
Multiplexed cystic fibrosis reference material (34 variants) | Seracare CF-Extended (#2040-0009)
MLH1/MSH2 Exon Copy Number Reference Panel | https://www.seracare.com/Controls---Reference-Materials-NGS-Inherited-Disease/
American Type Culture Collection (ATCC) | atcc.org
European Collection of Authenticated Cell Cultures (ECACC) | https://www.phe-culturecollections.org.uk/collections/ecacc.aspx
German Collection of Microorganisms and Cell Cultures | https://www.dsmz.de/
Japanese Collection of Research Bioresources | https://cellbank.nibiohn.go.jp/english/
Riken Bioresource Center | https://mus.brc.riken.jp/en/
Institute for Reference Materials and Measurements | https://ec.europa.eu/jrc/en/reference-materials
National Genetics Reference Laboratory Manchester | http://www.ngrl.org.uk/
Medical Device Innovation Consortium (MDIC) SRS Landscape Analysis Report | SRS Landscape Analysis Report - MDIC
WHO International Biological Reference Preparations | HERE