## Before you begin
Before you begin: How to use this workbook
1. Start on Tab A, Assay Requirements
Aside from defining clinical content, the most important aspect of designing a new test is the setting of requirements. This will inform all further work and decision making during the process. Requirements such as turnaround time (TAT), sample material to be validated (eg, saliva, blood), chemistry or technology, cost of goods for the final assay, and needs for future scaling etc. should be considered. The table in Tab A contains suggestions for requirements, but there are likely many other aspects that are unique to each test or laboratory. Requirement setting meetings should include all departments, and all decisions should be carefully discussed, documented, and signed off by all stakeholders.
2. Continue to Tab B, Initial Assay Configuration
This phase translates the requirements to an initial assay design. The target calculator estimates the size of the genomic region to be covered, and given a particular sequencing instrument/kit combination, how many samples can be sequenced per run and how many samples can be comfortably processed per batch. Note that there are often several kits and/or flowcell sizes available for each sequencing instrument, with varying data outputs, costs, and run times. If the assay design involves purchasing a new sequencing instrument, careful consideration must be given to each of these factors as well as requirements for future scaling. Since these offerings change regularly, first contact instrument vendors and fill in the sequencing instrument table to assess the current possibilities and limitations. Another important consideration is the assay methodology, ie, amplicon vs capture based. Although there are many nuances to this choice, in general more genes can be interrogated at the same time using a capture-based assay. Amplicon technology is very useful for technically challenging regions where the specificity inherent in careful primer design or the flexibility offered by PCR optimization possibilities is required (eg, to address GC rich regions). There are limits to the number of amplicon primers that can be used in a single tube multiplex, but note that various commercial offerings extend these limits and improve workflow, for example by using fluidics to mix primers and templates in separate chambers to avoid unfavorable primer/primer interactions. Capture-based technologies offer more flexibility in terms of parallelizing multiple tests both during library processing and sequencing, as well as for adding future content, but have limited optimization possibilities. As a rule, if more than 15 genes are planned for a test, and nothing suggests that the regions require special interventions for specificity, a capture-based assay may be the best choice. Choose both the sequencing instrument and the methology after careful consideration of all the requirements on tab A.
Specimen identity should be tracked in order to detect sample mix ups. Confirmatory testing of reported variants is accepted practice, but note that this will not cover negative test results. To extend the identity check to all samples, more extensive QC such as testing a predetermined set of common variants (fingerprinting) may be performed.
Ensure that the completed Initial Assay Design will theoretically fulfill all "must have" Assay Requirements from Tab A. If not, reassess at-risk requirements or initial assay design.
3. Continue to Tabs C and D, Pilot Run & Iteration
In TAB C, list the samples you will use to assess the performance of the assay in the first test run. This run is intended to confirm that the assay either already fullfills the requirements or needs optimization to reach fulfillment. Be conservative when choosing the number of samples to run, and consider running a number of samples in duplicate with different barcodes in order to be able to pool the data from identical samples to simulate a decrease in the sample number and concurrent increase in sequencing coverage. Choose a limited number of samples with relevant features as defined during test content design to include all variant types the assay is intended to detect, but reserve the majority of samples for validation. Prepare the draft SOP before processing any samples, and while performing the test run, take notes on workflow observations and possibilities for future optimization. It is important to process the samples in a manner as close to the envisioned final assay as possible (ie, without extensive pauses or slow pipetting), to ensure the process is reliably reproducible. Once the run is complete, analyzing the data in multiple ways is recommended (ie, using different variant callers) before assessing concordance with known samples.
Tab D provides a framework for estimating the operational impact of suboptimal read depth across the test region of interest
4. Analyze Data & Iterate as Necessary
After performing the first test run, carefully assess quality metrics against the Assay Requirements in Tab A. If any of the requirements are not met, consider how the assay can be improved or reconfigured to reach them. Alternatively, additional research may be needed to decide if a given requirement is truly necessary. Assay adjustments may require additional test runs to sufficiently optimize the assay before entering into the validation phase.
5. Define Acceptance and Rejection Criteria (Pass/Fail Metrics), Tab E
Define acceptance and rejection criteria for the results generated by the wet bench analytical process (MOL.36010). Criteria must be based on metrics and quality control parameters established during test optimization and utilized during validation. List the criteria on Tab D: Pass Fail QC. At this point, your SOP should be finalized and you can begin thorough validation runs.
© 2024 College of American Pathologists (CAP). All rights reserved.
CLSI. Human Genetic and Genomic Testing Using Traditional and High-Throughput Nucleic Acid Sequencing Methods. 3rd ed. CLSI guideline MM09. Clinical and Laboratory Standards Institute; 2023.
[

## A_Assay_Requirements
A. Assay Requirements & Test Element Specifications
Please note that this checklist is designed to provide a minimum list of critical considerations and metrics. These need to be carefully reviewed and modified to meet specific test development needs. Laboratories should cross reference this worksheet with additional worksheets as applicable.
1. Operational considerations | Comments/Notes | Requirements met by proposed technology? (Fill in as you go through the test optimization process (subsequent tabs))
Required/expected TAT? (Rapid considerations? Factor in time for repeats, confirmation, and interpretation)
Single sample ad hoc testing or sample batching?
Ability to combine different tests (existing or planned) within a sequencing run?
Expected volume and scaling projections?
2. External forces and internal limitations | Comments/Notes | Requirements met by proposed technology? (Fill in as you go through the test optimization process (subsequent tabs))
Specific market expectations that need to be met?
Restrictions due to existing instrumentation?
Additional laboratory equipment needed?
Limitations due to informatics infrastructure (including data storage)
Maximum possible cost of goods (all inclusive)?
3. Method Selection | Comments/Notes | Requirements met by proposed technology? (Fill in as you go through the test optimization process (subsequent tabs))
Target selection method (if applicable): Amplicon-based, capture-based, etc. - See notes in "before you begin"
Single-end/pair-end/mate-pair sequencing?
Multiplexing & molecular barcoding strategy if applicable, document avoiding barcode collision, mis-identification, mis-sorting and barcode hopping
Sample tracking methods - e.g. ID panel, plasmid spike-ins, LIMS custody
Library preparation methods - technologies and kits, document limitations and critical factors
QC methods do detect contamination
Considerations around library complexity (e.g., how much phiX should be loaded along with sequencing libraries)
4. Target Performance Requirements | Comments/Notes | Requirements met by proposed technology? (Fill in as you go through the test optimization process (subsequent tabs))
Target AVERAGE read depth across interpreted region? (e.g. 150x) The number of unique reads on average for each given nucleotide within your region of interest
Target per base MINIMUM read depth (i.e., how deeply does each position need to be covered for accurate variant calling? This will depend on variant types, typically indels have slightly lower variant allele fractions which might necessitate higher coverage for accurate calling
Target Total % Base-Pair completeness (e.g. 95% @20x) The percent of bases within your region of interest that are covered at the target minimum read depth
Which genes (if any) have to be "base-pair complete"? E.g., does every base have to meet the minimum read depth? If yes, then must anticipate and plan for "fill-in", e.g. by Sanger sequencing; if no, then must plan to calculate and indicate the fraction of the gene not covered on the clinical report.
Target Analytical Performance Metrics (see also worksheet "TEST VALIDATION" for terminology and definitions) Any industry thresholds or standard of care to meet or exceed?
5. Supplemental Methods Considerations | Comments/Notes | Requirements met by proposed technology?
Difficult regions (homology, GC content)
CNV Detection
Other structural variant detection
Fill-in (Sanger) sequencing
Confirmatory methods - e.g., Sanger, ddPCR

## B_Initial_Assay_Configuration
B. Initial Assay Configuration
1. Sequencing instrument table: obtain information from vendors and literature to assess options (consider single-end/pair-end/mate-pair sequencing)
Vendor | Instrument | Kit | Output | Time | Cost Instrument | Cost Service contract | Cost Kit | Limitations
Illumina | MiniSeq | e.g., Mid Out-Put (2 x 150bp) | 2.1–2.4 Gb | 17 hrs |  |  |  | repeat variations etc may not be adequately covered by short read technologies
Illumina | MiSeq | e.g., v2 (2 x 150bp)
Illumina | MiSeq | e.g., v3 (2 x 300bp)
Illumina | NextSeq 550 | High Output
Illumina | NextSeq 550 | Mid Output
Illumina | NextSeq 1000 | [6 empty] | Vendor only allows for use with whole genome sequencing.
Illumina | NovaSeq 6000 | [6 empty] | A very large number of samples must be pooled to be cost effective.
Illumina | NovaSeq X
Thermo Fisher | Ion GeneStudio S5
Thermo Fisher | PGM
Element Biosciences | AVITI
Pacific Biosciences | Onso
etc.
2. Target calculator: Determine your target region size
| Number of Exons | Size of Coding Region |  | Notes
Calculate from worksheet 1_Test_Content_Design | Total number of exons | Total Size of Target |  | Refer to Sheet 1 Test Content and bring in calculated content size
For capture-based NGS tests Buffer zone outside each exon (50 bp suggested) | Total number of exons x 2 x 50 bp | #VALUE! |  | A buffer zone allows for high quality coverage of the edges of the region of interest (eg, canonical splice sites)
Total target size (bp) | Total size of target + buffer zone | #VALUE! |  | This will inform which instrument/kit combination may need to be used
3. Estimate how many samples can fit on a run
Output of sequencing run (bases) |  |  |  | Output from the choosen instrument and kit from the sequencing instrument table above.
Desired average coverage (X) |  |  |  | Choose the desired average coverage (100X is a good starting point for germline tests, but expect to refine after performing the test run)
Theoretical # of samples per run | Output/bases needed | #VALUE!
Expected # of samples per run | Theoretical # of samples/run divided by 2 | #VALUE! |  | A good rule of thumb is to decrease by 50% from the theoretical number of samples to accommodate uneven coverage, pooling, and suboptimal sequencing runs.
Expected # of runs per month
Expected monthly throughput | Expected # of samples per run * runs per month | #VALUE!
4. Define confirmatory/sample identity testing approach
Notes
Are Variants confirmed prior to reporting? | Yes/No
Assay to be used
Is sample identity tracked for all samples? | Yes/No
Assay to be used
5. Make initial cost estimate per sample
Notes
Run cost divided by expected number of samples/run
Cost of extraction
Library preparation costs
Additional consumables
Confirmatory and/or supplemental assay costs
Labor estimate
Overhead estimate (incl. instrument depreciation)
Other
Total cost per sample | Total of above | 0
Total cost with 10% rework rate | Total cost * 1.1 | 0 |  | Start with a conservative rework rate as specimens often need additional processing
6. Design review: compare against requirements
Notes
Check that all requirements from Tab A have been met and addressed
Cost reduction options include: different sequencer/kit, different library preparation kit or method, eliminating "nice-to-have" target regions etc.
If the cost is higher than the requirement but reasonably close to meeting expectations, bear in mind that several conservative assumptions were made. It is likely worth proceeding with the test run to assess actual vs expected metrics

## C_Pilot_Run_&_Iteration
C. Pilot Runs & Iteration
1. Quantitation Requirements - Additional Test Element Specifications
Volume, quality, and concentration needed (consider excess for failures, repeats, confirmation)
Quantification specificity and methods (dsDNA, ssDNA, fluorescence, UV, Spectrometry)
2. Pilot Runs: Target samples which will best help answer performance metric questions
Note this table can also be found in the Test Validation Worksheet.
This example reflects a validation scheme designed for the MiSeq: 10 samples/lane
Concept: Base-line validation satisfied by NA12878 (or other well characterized reference sample(s)) Fill run up with other reference materials depending on gene-/disease-specific needs
SAMPLE ID | RUN ID | WITHIN-RUN ID | SPECIMEN TYPE | DESCRIPTION | PURPOSE | PURPOSE (DETAIL) | SPECIFIC VARIANT/S OF INTEREST | VARIANT TYPE | MEASURED METRIC
VAL-1 | Run-1 | Sample-1 |  | NA12878 | Base-Line Validation | Sens/PPV/Cov/Completeness | N/A | N/A | Intra-run precision
VAL-2 | Run-1 | Sample-2 |  | NA12878 | Base-Line Validation | Sens/PPV/Cov/Completeness | N/A | N/A | Intra-run precision
VAL-3 | Run-1 | Sample-3 |  | NA12878 | Base-Line Validation | Sens/PPV/Cov/Completeness | N/A | N/A | Intra-run vprecision
VAL-4 | Run-1 | Sample-4 |  | Anonymized patient specimen | Base-Line Validation | Variant Type Prevalent in Gene | List Variant Here | IN/DEL | Sensitivity
VAL-5 | Run-1 | Sample-5 |  | Anonymized patient specimen | Base-Line Validation | Variant Type Prevalent in Gene | List Variant Here | CNV | Sensitivity
VAL-6 | Run-1 | Sample-6 |  | Anonymized patient specimen | Gene-Specific Validation | Mosaicism (Variant at 20%) | List Variant Here | SNV | Limit of Detection
VAL-7 | Run-1 | Sample-7 |  | Anonymized patient specimen
VAL-8 | Run-1 | Sample-8 |  | Anonymized patient specimen
VAL-9 | Run-1 | Sample-9 |  | Anonymized patient specimen
VAL-10 | Run-1 | Sample-10 |  | Anonymized patient specimen
3. Sample Data Analyses
Instrument Run Level Metrics (Determine range, mean, eveness, and identify outliers. (Example metrics, select what is most meaningful in your laboratory. Range and average will not apply if only one run or lane was submitted.)
| QC Metrics | Average | Low End | High End | Interim Threshold
Density (K/mm2)
Reads
Reads PF
% Reads PF
Error Rate
% >= Q30
% Aligned
Phasing
Prephasing
Per Sample General Performance Metrics (Determine range, mean, evenness, and identify outliers (Example metrics, select what is most meaningful in your laboratory)
| QC Metric | Average | Low End | High End | Interim Threshold
% of pool
# Reads
% Reads PF
% of raw clusters per lane
% Perfect Index Reads
% of >= Q30 Bases (PF)
Average coverage across defined ROI
Gene/Exon Level Coverage
Completeness (% of desired bases covered with X reads)
GC Drop
AT Drop
Library complexity & % Duplication rate
• Is there excess per sample coverage? Can more samples be pooled?
• Is there too little per sample coverage? Use less samples for pooling?
• Any critical gaps in coverage? Gene or exon level
• Any critical gaps in coverage? Variant level
• Consider dropping low priority regions with poor coverage
Developing Variant Calling Thresholds
• Use reference materials with known variants to optimize bioinformatic thresholds for variant calling (QD, FS, AX, MQ, etc.).
Analyses to Determine Additional Optimization Needs
• Analytical performance metrics (analytical sensitivity, specificty, etc) - see Test Validation Worksheet for a complete list
• Do all specimen types perform adequately?
• Are there any new or unanticipated limitations after pilot runs?

## D_completeness of coverage asse
Completeness of coverage (read depth) considerations
• For some tests (especially diagnostic gene panels), completeness of coverage can be critical. It is customary to fill in undercovered regions (e.g. by Sanger sequencing) to ensure optimal analytical sensitivity. • It is critical to calculate completeness of coverage over the test region. This exercise is helpful in estimating the operational impact of "fill in sequencing" (on cost and turnaround time) ahead of launch and make adjustments if needed. • Note: This is especially important when the assay is larger than the test region of interest (eg: WES or WGS assay run but reported content is a diagnostic gene panel). Exome/genome wide average coverage metrics can be misleading as they can mask critical coverage gaps in clinically important genes where fill-in sequencing needs to be performed.
EDUCATIONAL EXAMPLE
Step 1: Use Region of Interest (ROI) from Worksheet "Test Content Design - TAB C_Target_Region_Technical_Info_V3.0"? Scenario: For each interval and each tested run condition: What % of samples FAILED (most conservative threshold: >1 base >min read depth)
From Test Content Design Worksheet: |  |  |  |  | run condition 1 | run condition 2 | run condition 3
Gene ID | Chr | Start | End | Size | % of samples FAILED | % of samples FAILED | % of samples FAILED
Highlighted cells: >5% samples in run FAILED
ATM | 11 | 108126922 | 108127081 | 160 | 0% | 0% | 0%
ATM | 11 | 108164025 | 108164217 | 193 | 50% | 4.167% | 4.167%
ATM | 11 | 108216457 | 108216648 | 192 | 29.17% | 25% | 8.333%
BAP1 | 3 | 52443845 | 52444037 | 193 | 70.83% | 16.67% | 0%
BARD1 | 2 | 215657008 | 215657182 | 175 | 45.83% | 16.67% | 8.333%
BLM | 15 | 91337384 | 91337600 | 217 | 75% | 75% | 8.333%
BRCA1 | 17 | 41205886 | 41205899 | 14 | 54.17% | 8.333% | 4.167%
BRCA1 | 17 | 41219277 | 41219290 | 14 | 25% | 20.83% | 4.167%
DIS3L2 | 2 | 233194510 | 233194719 | 210 | 8.333% | 0% | 0%
DKC1 | X | 154001393 | 154001537 | 145 | 0% | 8.333% | 8.333%
FANCE | 6 | 35420309 | 35420583 | 275 | 91.67% | 16.67% | 4.167%
FANCM | 14 | 45623887 | 45624038 | 152 | 0% | 0% | 0%
GPC3 | X | 132795745 | 132795891 | 147 | 8.333% | 16.67% | 0%
MLH1 | 3 | 37049635 | 37049647 | 13 | 4.167% | 8.333% | 8.333%
MLH1 | 3 | 37067113 | 37067511 | 399 | 100% | 100% | 79.17%
MSH2 | 2 | 47641395 | 47641571 | 177 | 75% | 41.67% | 16.67%
MSH2 | 2 | 47656807 | 47656819 | 13 | 95.83% | 83.33% | 41.67%
MSH2 | 2 | 47672105 | 47672117 | 13 | 8.333% | 8.333% | 4.167%
MSH3 | 5 | 79950533 | 79950796 | 264 | 50% | 0% | 0%
PMS2 | 7 | 6029418 | 6029599 | 182 | 8.333% | 20.83% | 0%

## E_QC+Pass_Fail
D. Quality Control Steps and Pass/Fail Criteria
Example Test Workflow With QC Stops and Pass/Fail Criteria
ASSAY STEP | QC METHOD | PASS / FAIL CRITERION | ACTION IF FAILURE | NOTES
DNA Extraction | Qubit / Picogreen | Minimum 3ug and 24ng/ul | Re-extract remaining sample or request new sample | Input DNA quality and quantity
Fragmentation | TapeStation gel image- sizing assay | 150-200bp ± 10% with >1000 FU intensity | If FU intensity is <1000, it is recommended to re-quant sample and repeat shear to ensure sufficient Input into library construction | Measures size of fragments and distribution of sizes
Library construction & PCR | Qubit / Picogreen | Samples must have a concentration of 38ng/ul (total volume is 26.4ul) or higher in order to reach the 1ug needed as input for hybridization | If the total amount of sample is <1ug, repeat PCR using the remaining Pre-PCR library | Output library quality and quantity
Hybridization and target selection QC1 | Qubit / Picogreen | >15ng total per sample. concentration >3ng/ul | Discuss with supervisor, based on predicted final pool of batch | Hybridization quantity per sample
Hybridization and target selection QC2 | TapeStation gel image- sizing assay | Fragments ranging from 250-350bp | Discuss with supervisor | Measures size of fragments and distribution of sizes
Final Pool | Qubit / Picogreen | 1.0nM of the target 14nM (13nM – 15nM). | Discuss with supervisor | Measures concentration of final pool for use on sequencing instrument and cluster generation
General sequencing performance | Cluster Density | 220-240 k/mm2 | Re-launch sequencer with adjusted loading amount | Measures efficient use of Flowcell maximize data, minimize overclustering
General sequencing performance | Completeness Across Test Region | 95% bases @ 20x | Discuss with supervisor | Gives a sense of how much of test region is uncovered
General sequencing performance | Average Coverage | >=150X | Discuss with supervisor
General sequencing performance | # Sanger Sequencing Fill-Ins Across Panel ROI | <10 Sanger Amplicons | Discuss with supervisor | Cost and TAT increases as Sanger Fill-Ins increase, at some point might be cheaper to re-run library construction from scratch