## Before you begin
Before you begin: How to use this workbook
1. Designing Your Validation Study
Analytical validation in a clinical laboratorty is carried out in a "beginning-to-end" manner, ie, covering ALL steps included in the test. Traditionally, this includes all steps from nucleic acid extraction to final result (clinical report). However, currently parts of the NGS process, such as sequencing, bioinformatics, or clinical interpretation may be offered separately by other clinical entities. The validation requirements for these partial services are different from those applied to a full service laboratory and are described separately below.
Please note: Validations for supplemental assays may not be covered by the illustrated NGS validation setup in subsequent tabs, but consider using a similar framework.
The following hybrid validation design is recommended for laboratory-developed tests (LDTs):
A) Baseline analytical performance:
Establish general analytic performance by variant type that must be detected by the test (ie, SNVs, indels, CNV, other) and each specimen type the test must accommodate (eg, whole blood, saliva, amniotic fluid, etc.).
For variants other than SNVs, establish performance for different categories (eg, different size ranges for indels or CNVs).
Maximize # of variants tested for increased statistical confidence.
Note: This component is "methods-based" (Aziz et al. 2015, [checklist codes]), ie, performance data from multiple tests (eg, multiple gene panels) can be added to enhance statistical power provided that the test protocols are identical.
Commonly used samples for this type of validation analysis are highly characterized samples, such as those characterized by the Genome-in-a-Bottle Consortium (http://jimb.stanford.edu/giab/; Zook 2016 PMID 27271295), or CDC Genetic Testing Reference Materials Coordination Program (GeT-RM; https://wwwn.cdc.gov/clia/Resources/GetRM/). See "Test Content Design" worksheet for more information on reference materials.
B) Additional studies as needed. Examples include but are not limited to:
A medically critical gene with a high prevalence of a variant type where the baseline performance does NOT provide sufficient statistical confidence (eg, the number of variants for a certain variant type that can be derived from the reference sample(s) is too low). Example: CNV detection for Duchenne Muscular Dystrophy (DMD).
Common pathogenic variant(s) exist and baseline performance does NOT provide sufficient confidence for the type of variant. Example: the common, recurring 3 base deletion in the CFTR gene (c.1521_1523delCTT; p.F508del).
Gene or clinically relevant variant falls within a difficult to sequence region, which may have different analytical performance metrics than general regions analyzed in a methods-based validation. Example: Genes with regions of high homology to other regions of the genome (e.g.PMS2, STRC, Mandelker 2016).
2. Start on Tab A, Performance Metrics
Review performance metrics that need to be recorded.
Tab A_Performance_Metrics provides a comprehensive summary of analytical performance metrics that are typically assessed during validation. Note that terminology and definitions is not entirely harmonized across different publications/recommendations that have been issued to date.
NGS testing that is designed to detect variants represents a qualitative test where the signal obtained is translated to a limited number or predefined outcomes (homozygous reference, heterozygous reference/variant, homozygous variant).
While there are many acceptable approaches to satisfy the commonly used analytical performance metrics, keep in mind that the following are the core principles that are addressed:
1. How well does the assay detect variants that are known to be present?
2. How high is the likelihood for detecting variants that are not present (false positive results)?
3. How "robust" is the assay when certain conditions are altered (eg, a different sequencing instrument is used)?
3. Continue to Tabs B and C (Example Validation Grid and Example Sample List)
Assemble your validation samples and design validation runs.
Please note that these tabs illustrate ONE possible approach that illustrates how a validation can be carried out.
4. Continue to Tabs D-G (Example Performance Metrics and Assay Limitations)
After analyzing the validation runs (see examples in Tabs D and E) and calculating the analytical performance metrics (see an example in Tab "F_Perf.Summary.Report"), review the achieved performance and ensure that it meets requirements for the intended use. If these requriements are NOT met, go back to the Assay Design & Optimization Worksheet. Otherwise, the test can be launched. It is advisable to summarize limitations on the report (see tab "G_Limitations+Clinical.Report.Example" for a representative example).
© 2024 College of American Pathologists (CAP). All rights reserved.
CLSI. Human Genetic and Genomic Testing Using Traditional and High-Throughput Nucleic Acid Sequencing Methods. 3rd ed. CLSI guideline MM09. Clinical and Laboratory Standards Institute; 2023.

## A_Performance_Metrics
A. Analytical Performance Metrics
See the CLSI Harmonized Terminology Database for a compilation of internationally accepted terminology: https://clsi.org/standards-development/harmonized-terminology-database/
ANALYTICAL PERFORMANCE METRIC |  | FORMULA/ DEFINITION | COMMON SENSE TRANSLATION | ADDITIONAL COMMENTS | EXAMPLE APPROACHES | LABORATORY ACCREDITATION REQUIREMENTS
LDT terminology | IVD terminology |  |  |  |  | CAP CHECKLIST
A) Variant detection
Analytical Sensitivity | Positive Percent Agreement | TP/[TP+FN] | The fraction of variants known to exist by a designated gold standard that are correctly identified as such by the new test. Notes: • Sometimes, the term "analytical sensitivity" is used to express "lower limit of detection" (common in somatic sequencing applications as well as IVD applications). • Bioinformaticians commonly use "recall" instead of "analytical sensitivity" (PMID: 30858580) | • Should be calculated separately for all types of variants relevant to the test/disease. • Confidence intervals should be calculated. While there is no mandatory degree of confidence, a widely utilized minimum value is 95%. The LOWER bound of the confidence interval is the one that needs to meet target performance. Sample confidence interval calculator: http://www.pedro.org.au/english/downloads/confidence-interval-calculator/. • Values for separate tests (i.e. different groups of genes) can be added to derive a cumulative value provided that assay methodology is the same. • For sequencing tests that interrogate a large number of bases, it is important to calculate the Technical PPV (TP/TP+FP), Note that this metric is not part of the traditional CLIA performance metrics. Calculating only analytical specificity (where "TN" represents the number of WT bases correctly identified as WT), a high false positive rate can be masked by the usually very large number of WT bases in the denominator of the formula. | • Analytical validation of sequencing tests that are intended to detect novel variants require two validation layers: (a) methods-based (maximize number of variants, pathogenicity does not matter) and (b) gene/disease specific (must include reference mateirals for common, pathogenic variants). • Use well characterized reference samples (e.g. NA12878, Coriell samples) • Use previously tested, clinical specimens • Alternative reference materials (e.g. for rare but essential variants) such as engineered cell lines, linearized plasmids or other artificial constructs diluted in genomic DNA • In silico mutagenized files may be used to supplement (but not supplant) patient samples and physical specimens, and these can be especially helpful in generating difficult to source variants (e.g., indels of varying sizes) | MOL.31130
Analytical Specificity | Negative Percent Agreement | TN/[TN+FP] | The test's ability to accurately call WT (reference) bases (to obtain negative results in concordance with negative results obtained by the reference method. Notes: • Sometimes, term "analytical specificity" is used to express "interference" and "cross-reactivity" |  |  | MOL.31130
Technical Positive Predictive Value | Technical Positive Predictive Value | TP/[TP+FP] | Likelihood that a variant call is a true positive. Notes: • Bioinformaticians commonly use "precision" instead of "positive predictive value" (PMID: 30858580)
Accuracy | Accuracy | [TP+TN]/ [TP+FP+TN+FN] | Accuracy can be equated to sensitivity + specificity. Note that this metric is not commonly used (Jennings 2009 and Mattocks 2010). |  |  | MOL.31130
False Negative Rate | usually not required | 1 - Sensitivity | The false negative rate is the proportion of the individuals with a known positive condition for which the test result is negative. This rate is sometimes called the miss rate. |  |  | MOL.31590
False Positive Rate | usually not required | 1 - Specificity | The false positive rate is the proportion of the individuals with a known negative condition for which the test result is positive |  |  | MOL.31590
B) Performance across altered conditions and equipment
Precision | Precision | N/A | Closeness of agreement between independent test results obtained under stipulated conditions (a measure of random error). | Precision is usually divided into two sub-categories: 1. repeatability (within-run precision) 2. reproducibility (conditions are varied) | • Performance can be measured using a variety of QC metrics established by the laboratory. Examples can include (not limited to): fragment size and distribution, average coverage, % reads on target. • If using reference specimen, such as NA12878, analytical sensitivity and specificity can also be used as measures (refer to Example Validation Grid). | MOL.31130
Precision - Repeatability | Precision - Repeatability | Intra-Run Precision | Degree to which the same sequence is derived when sequencing a reference sample multiple times, under the same conditions. |  |  | MOL.31130
Precision - Reproducibility | Precision - Reproducibility | Conditions Altered | Closeness of agreement between repeated tests for the same sample under changed conditions (instrument, laboratory, operator, sample source, reagent lot, run, day, time of day). |  |  | MOL.31130
Robustness | Robustness /Guard-banding | N/A | Test precision given small, deliberate changes in test conditions (e.g. variations in storage conditions, sample concentration, etc).
C) Other metrics
Limit of Detection | Limit of Detection | N/A | Lowest amount/concentration of analyte that gives an informative result at a given level of confidence for each variant type. | This metric also applies to the following: a) mosaicism (lowest percent detectable) b) mitochondrial heteroplasmy (lowest percent detectable)* |  | MOL.31130 MOL.36153
--- | Interference | N/A | –– | The effect of substances in the sample that cause the test to give wrong results. |  | MOL.31130
--- | Cross-Reactivity | N/A | The ability of a drug, metabolite, a structurally similar compound other than the primary measurand, or even an unrelated compound, to affect the measurement procedure (per CLSI C52, C63) Reactivity of a test reagent with a substance in the sample other than the analyte tested. | In molecular testing, this term is usually used to describe cross-reaction with homologous sequences, often leading to a false positive result |  | MOL.31130
Reportable Range | Reportable Range | N/A | The reportable range encompasses the full range of reported values. | For qualitative tests that would include all reportable outcomes (e.g. homozygous wild type, heterozygous or homozygous variant). |  | MOL.31130
Reference Interval | Reference Interval | N/A | The reference interval is the range of results expected in the normal population. | For NGS, this refers to the normal variation of sequences within the population that the assay is designed to detect. |  | MOL.31255
TP = True Positive | FP = False Positive | TN = True Negative | FN = False Negative
NOTES
* Mitochondrial sequencing is not covered by MM09

## B_Example_Validation_Grid
B. Sample Validation Run Designs
NOTE: The examples below are designed to be a minimal validation to assess all required performance metrics. Additional requirements may exist depending on which regulatory agency a validation is approved by (such as for IVDs that are submitted to the US Food and Drug Administration (FDA))
Example 1: HiSeq 2500 (8 lanes) 10 samples/capture, run 3 times (same run) + 2 times on 2 additional runs | [10 empty] | Example 2: MiSeq (1 lane) 10 samples/capture
load same pool 3 times | RUN 1 - LANE 1 | LANE 2 | LANE 3 | LANE 4 | LANE 5 | LANE 6 | LANE 7 | ◄--- | INTRA-RUN Run same pool of 10 three times on same flow cell (assesses instrument variability) | ---► | RUN 1 - LANE 1 | NOTE Intra-run variability limited to NA12878
| NA12878 - A - barcode 1 | NA12878 - A - barcode 1 | NA12878 - A - barcode 1 | OPTIONAL LANES, MAY HAVE NEED TO VALIDATE ADDITIONAL SAMPLES 1. Samples with frequent pathogenic variants 2. Samples carrying variant type/s of interest (e.g. In/Dels, CNVs, structural variants) | [6 empty] | NA12878 - A - barcode 1
| NA12878 - A - barcode 2 | NA12878 - A - barcode 2 | NA12878 - A - barcode 2 | [7 empty] | NA12878 - A - barcode 2
| NA12878 - A - barcode 3 | NA12878 - A - barcode 3 | NA12878 - A - barcode 3 | [7 empty] | NA19240 - A
| NA19240 - A | NA19240 - A | NA19240 - A | [7 empty] | Sample 4
| Sample 5 | Sample 5 | Sample 5 | [7 empty] | Sample 5 | NOTE Samples 4-10 may differ between runs based on need to add disease specific samples
| Sample 6 | Sample 6 | Sample 6 | [7 empty] | Sample 6
| Sample 7 | Sample 7 | Sample 7 | [7 empty] | Sample 7
| Sample 8 | Sample 8 | Sample 8 | [7 empty] | Sample 8
| Sample 9 | Sample 9 | Sample 9 | [7 empty] | Sample 9
| Sample 10 | Sample 10 | Sample 10 | [7 empty] | Sample 10
Same samples - different library preps | RUN 2 - LANE 1 | [9 empty] | RUN 2 - LANE 1
| NA12878 - B - barcode 1 | [9 empty] | NA12878 - B - barcode 1
| NA12878 - B - barcode 2 | [9 empty] | NA12878 - B - barcode 2
| NA12878 - B - barcode 3 | [9 empty] | NA19240 - B
| NA19240 - B | [9 empty] | Sample 4
| Sample 5 | [9 empty] | Sample 5
| Sample 6 | [9 empty] | Sample 6
| Sample 7 | [9 empty] | Sample 7
| Sample 8 | [9 empty] | Sample 8
| Sample 9 | [9 empty] | Sample 9
| Sample 10 | [9 empty] | Sample 10
| RUN 3 - LANE 1 | [9 empty] | RUN 3 - LANE 1
| NA12878 - C - barcode 1 | [9 empty] | NA12878 - C - barcode 1
| NA12878 - C - barcode 2 | [9 empty] | NA12878 - C - barcode 2
| NA12878 - C - barcode 3 | [9 empty] | NA19240 - C
| NA19240 - C | [9 empty] | Sample 4
| Sample 5 | [9 empty] | Sample 5
| Sample 6 | [9 empty] | Sample 6
| Sample 7 | [9 empty] | Sample 7
| Sample 8 | [9 empty] | Sample 8
| Sample 9 | [9 empty] | Sample 9
| Sample 10 | [9 empty] | Sample 10
| ▲ I | [9 empty] | ▲ I
| INTER-RUN Set up 10 samples 3 times from start to finish. This reflects clinical reality of needing to re-run samples | [9 empty] | INTER-RUN Set up 10 samples 3 times from start to finish. This reflects clinical reality of needing to re-run samples

## C_Example_Sample_List
C. Example Sample List
This example reflects a validation scheme designed for the MiSeq: 10 samples/lane, 3 replicate runs. Concept: Methods-based validation satisfied by NA12878 (or other well characterized reference sample/s). Fill run up with other reference materials depending on gene/disease specific needs
Sample ID | Run ID | Within Run ID | Description | Specimen Type | Purpose | Purpose (Detail) | Specific Variant(s) of Interest | Variant Type | Measured Metric
VAL-1 | Run-1 | Sample-1 | NA12878 | Cell Line | Methods-based Validation | N/A | N/A | N/A | Intra-run precision
VAL-2 | Run-1 | Sample-2 | NA12878 | Cell Line | Methods-based Validation | N/A | N/A | N/A | Intra-run precision
VAL-3 | Run-1 | Sample-3 | NA12878 | Cell Line | Methods-based Validation | N/A | N/A | N/A | Intra-run vprecision
VAL-4 | Run-1 | Sample-4 | Anonymized patient specimen | Blood | Methods-based Validation | Variant Type Prevalent in Gene | List Variant Here | IN/DEL | Sensitivity
VAL-5 | Run-1 | Sample-5 | Anonymized patient specimen | Blood | Methods-based Validation | Variant Type Prevalent in Gene | List Variant Here | CNV | Sensitivity
VAL-6 | Run-1 | Sample-6 | Anonymized patient specimen | Blood | Gene-Specific Validation | Mosaicism (Variant at 20%) | List Variant Here | SNV | Limit of Detection
VAL-7 | Run-1 | Sample-7 | Anonymized patient specimen | Blood
VAL-8 | Run-1 | Sample-8 | Anonymized patient specimen | Blood
VAL-9 | Run-1 | Sample-9 | Anonymized patient specimen | Blood
VAL-10 | Run-1 | Sample-10 | Anonymized patient specimen | Blood
VAL-11 | Run-2 | Sample-1 | NA12878 | Cell Line | Methods-based Validation | N/A | N/A | N/A | Inter-run variability
VAL-12 | Run-2 | Sample-2 | Anonymized patient specimen | Saliva
VAL-13 | Run-2 | Sample-3 | Anonymized patient specimen | Saliva
VAL-14 | Run-2 | Sample-4 | Anonymized patient specimen | Saliva
VAL-15 | Run-2 | Sample-5 | Anonymized patient specimen | Saliva
VAL-16 | Run-2 | Sample-6 | Anonymized patient specimen | Saliva
VAL-17 | Run-2 | Sample-7 | Anonymized patient specimen | Saliva
VAL-18 | Run-2 | Sample-8 | Anonymized patient specimen | Saliva
VAL-19 | Run-2 | Sample-9 | Anonymized patient specimen | Saliva
VAL-20 | Run-2 | Sample-10 | Anonymized patient specimen | Saliva
VAL-21 | Run-3 | Sample-1 | NA12878 | Cell Line | Methods-based Validation | N/A | N/A | N/A | Inter-run variability
VAL-22 | Run-3 | Sample-2 | Anonymized patient specimen | Buccal Swab
VAL-23 | Run-3 | Sample-3 | Anonymized patient specimen | Buccal Swab
VAL-24 | Run-3 | Sample-4 | Anonymized patient specimen | Buccal Swab
VAL-25 | Run-3 | Sample-5 | Anonymized patient specimen | Buccal Swab
VAL-26 | Run-3 | Sample-6 | Anonymized patient specimen | Buccal Swab
VAL-27 | Run-3 | Sample-7 | Anonymized patient specimen | Buccal Swab
VAL-28 | Run-3 | Sample-8 | Anonymized patient specimen | Buccal Swab
VAL-29 | Run-3 | Sample-9 | Anonymized patient specimen | Buccal Swab
VAL-30 | Run-3 | Sample-10 | Anonymized patient specimen | Buccal Swab

## D_Valid_Results_1_Gen_Perf.
D. Sample Validation Summary (1) - General Performance Metrics
1. Brief Summary of Validation Runs
a) NA12878 run 4 times (runs 1-4) [inter-run variability]
b) NA12878 run in duplicate (A/B) 3 times (runs 2-4) [intra-run variability]
c) 30 additional reference samples with gene/disease-specific variants [disease-specific sensitivity]
NOTE: This is a sample summary showing a set of commonly used general metrics (run and sample level) - there are no formal requirements around which metrics to use
2. Sequencing-Level Technical Performance
Metric | Average | Min | Max | Description
Cluster Density (k/mm2) | 217.9167 | 152 | 312 | Density of clusters detected by image analysis
Clusters PF (%) | 80.48 | 60.24 | 88.7 | Percentage of clusters passing filtering
Reads (M) | 141.3796 | 98.58 | 202.61 | The number of clusters (in millions)
Reads PF (M) | 111.7254 | 82.08 | 147.74 | The number of clusters (in millions) passing filtering
%>=Q30 | 91.2046 | 85.28 | 95.03 | Percentage of bases with a quality score of 30 or higher
Alignment (%) | 5.82 | 3.78 | 7.22 | The percentage that aligned to the PhiX genome
Error Rate (%) | 0.31667 | 0.17 | 0.61 | Calculated error rate, as determined by the PhiX alignment
3. Sample-Level Technical Performance
Metric | Average | Min | Max | Description
Pass-Filter Reads | 4542463.6 | 2823364 | 5897048 | The total number of PF reads in the SAM or BAM file examined
% Pass-Filter Unique Reads | 84.13% | 78.99% | 86.8% | The percentage of PF reads that are not marked as duplicates
% Pass-Filter Unique Reads Aligned | 96.57% | 94.03% | 97.36% | The percentage of PF unique reads that aligned to the reference sequence
% Selected Bases | 69.09% | 51.78% | 78.6% | The fraction of PF aligned bases located on or near a baited region
Mean Bait Coverage | 483.8077 | 301.7454 | 656.9493 | The mean coverage of all baits in the experiment
Mean Target Coverage | 553.5301 | 343.2739 | 749.7404 | The mean coverage of a target region
% Bases >= 2X | 99.75% | 99.66% | 99.81% | The fraction of all target bases achieving 2X or greater coverage
% Bases >= 10X | 99.65% | 99.41% | 99.75% | The fraction of all target bases achieving 10X or greater coverage
% Bases >= 20X | 99.55% | 98.89% | 99.73% | The fraction of all target bases achieving 20X or greater coverage
% Bases >= 30X | 99.42% | 98.19% | 99.71% | The fraction of all target bases achieving 30X or greater coverage
AT Dropout | 1.25 | 0.85 | 4.12 | How undercovered <= 50% GC regions are relative to the mean
GC Dropout | 0.83 | 0.77 | 2.38 | How undercovered >= 50% GC regions are relative to the mean

## E_Valid_Res._2_analytical.perf.
E. Validation Results (2) - Analytical Performance Metrics
1. Methods-Based Validation
Analytical performance metrics derived from: NIST v2.19 (NA12878) - 4 runs (sample run in duplicate on runs 2-4)
| NA12878-Run 1 | NA12878-Run 2A | NA12878-Run 2B | NA12878-Run 3A | NA12878-Run 3B | NA12878-Run 4A | NA12878-Run 4B
Total Bases | 256357 | 256357 | 256357 | 256357 | 256357 | 256357 | 256357
True Negatives (TN) | 256211 | 256210 | 256212 | 256209 | 256211 | 256209 | 256212
False Positives (FP) | 0 | 1 | 0 | 0 | 1 | 1 | 0
False Negatives (FN) | 1 | 1 | 0 | 3 | 0 | 2 | 0
True Positives (TP) | 144 | 144 | 145 | 142 | 145 | 143 | 145
SNVs | 142 | 141 | 142 | 140 | 142 | 141 | 142
In/Dels | 2 | 3 | 3 | 2 | 3 | 2 | 3
Sensitivity (TP/TP+FN) (All Variants) | 99.31% | 99.31% | 100% | 97.93% | 100% | 98.62% | 100%
95% Confidence Interval | 96.20-99.88% | 96.20-99.88% | 97.42-100.00% | 94.09-99.29% | 97.42-100.00% | 95.11-99.62% | 97.42-100.00%
Sensitivity (SNVs) | 100% | 99.3% | 100% | 98.59% | 100% | 99.3% | 100%
95% Confidence Interval | 97.37-100.00% | 96.12-99.88% | 97.37-100.00% | 95.01-99.61% | 97.37-100.00% | 96.12-99.88% | 97.37-100.00%
Sensitivity (In/Dels)**see below | 66.67% | 100% | 100% | 66.67% | 100% | 66.67% | 100%
95% Confidence Interval | 20.77-93.85% | 43.85-100.00% | 43.85-100.00% | 20.77-93.85% | 43.85-100.00% | 20.77-93.85% | 43.85-100.00%
Specificity (TN/TN+FP) | 100% | 99.99% | 100% | 100% | 99.99% | 99.99% | 100%
95% Confidence Interval | 99.99-100.00% | 99.99-100.00% | 99.99-100.00% | 99.99-100.00% | 99.99-100.00% | 99.99-100.00% | 99.99-100.00%
Techn. Pos. Pred. Value (TP/TP+FP) | 100% | 99.31% | 100% | 100% | 99.32% | 99.31% | 100%
95% Confidence Interval | 97.42-100.00% | 96.20-99.88% | 97.42-100.00% | 97.42-100.00% | 96.22-99.88% | 96.17-99.88% | 97.42-100.00%
** The number of In/Dels present in the test region in this reference sample (NA12878) is too low to establish sufficient confidence - therefore, additional reference samples and/or synthetic reference materials need to be tested OR a limitation need to be stated in the clinical report methodology
2. Gene-/Disease-Specifc Validation . Example: Prevalent Pathogenic Variants
Test results from reference samples with prevalent pathogenic variants
Gene | Transcript | cDNA coordinates | Protein coordinates | Variant type | Variant Contribution | Detected? | Action
GENEA | NM_XXXXXX.Y | c.XXXX | p.YYYY | SNV | 5% of disease XYZ | YES
GENEB | NM_XXXXXX.Y | c.XXXX | p.YYYY | Duplication (20 bp) | 2% of disease XYZ | NO | Add limitation to report
GENEC | NM_XXXXXX.Y | c.XXXX | p.YYYY | SNV | 1% of disease XYZ | YES

## F_Perf.Summary.Report
F. Performance Summary Report
Considerations regarding non-SNV variants a) Small Insertions, Deletions, InDels, Duplications, etc: Ideally, those should all be assessed separately and broken down by size range. It is understood that this is currently impractical and infeasible due to the lack of apropirate reference materials (as these varaints are rare). At minimum, labs should analyze all types of variants in this category by size range. b) CNVs: The same applies here. These include deletions and duplications. Ideally, those are analyzed separately (as duplications pose greater challenges than deletions) c) Per CAP it is permissible to augment validations using in silico reference materials
Additional Performance Metrics
Metric | Formula / Definition | Test Performance*** (test ROI only) | Platform Performance*** (cumulative, see comment) | *** Test versus Platform (Methods based) calculations: For a methods-based validation, one can conbine several tests as long as they are run using identifcal procedures. For example, it is becoming common practice to run an exome or a genome and bioinformatically "slice" the test (eg a gene panel) from the data. In this case, one can use the entire exome or genome as a reference set.
Analytical Sensitivity (Substitutions) | Ability to detect all confirmed variants (TP/TP+FN) | 98.75% (158/160) | 98.28% (458/466) (CI95%: 96.65-99.13%)
Analytical Sensitivity (Insertions/Deletions, all sizes) |  | 73.33% (11/15) | 77.27% (17/22) (CI95%: 56.56-89.88%)
Analytical Sensitivity (Insertions/Deletions/etc, size range 1 - eg: 1-5 bp) |  | 80.00% (4/5) | 90.00% (9/10) (CI95%: 59.58-98.21%)
Analytical Sensitivity (Insertions/Deletions/etc, size range 2 - eg: 6-10 bp) |  | 100.00% (2/2) | 83.33% (5/6) (CI95%: 43.65-96.99%)
Analytical Sensitivity (Insertions/Deletions/etc, specify: size range 3 - eg 11+ bp) |  | 62.50% (5/8) | 50.00% (3/6) (CI95%: 18.76-81.24%)
Analytical Sensitivity (CNVs specify: size range 1 - eg: 1-5 exons)
Analytical Sensitivity (CNVs specify: size range 2 - eg: 5-10 exons)
Analytical Sensitivity (CNVs specify: size range 3 - eg: 10+ exons)
Analytical Specificity | Ability to correctly identify WT bases (TN/TN+FP)
Technical Positive Predictive Value (Substitutions) | Likelihood an NGS call is a true positive (TP/TP+FP)
Technical Positive Predictive Value (In/dels)
Technical Positive Predictive Value (CNVs)
Precision - Repeatability (Specify metric used to derive this - eg: analytical sensitivity) | Intra-Run Precision (Within-run variability of chosen metric)
Precision - Reproducibility (Specify metric used to derive this - eg: analytical sensitivity) | Conditions Altered EG: Between-run variability of chosen metric

## G_Limitations+Clinical.Report.E
Final Performance Checklist and Conclusions
Depending on the answers to the questions below, the test developer either needs to go back and continue optimizing the test or, if limitations are deemed acceptable, add those to the clinical report (example below).
1. Is the analytical sensitivity sufficiently high for the gene(s) and disease(s) of interest?
2. Are there existing benchmarks (eg, for well established genes) and if yes, were they met? Example 1: CFTR offered for diagnostic testing: This is a well established gene. The assay must be able to detect all common pathogenic CFTR variants. If the validation result shows suboptimal performance --> go back to further optimize performance, add supplemental assay, or list a limitation in the report methodology. Example 2: STRC offered for diagnostic testing: No benchmark exists due to technical difficulties associated with this gene (high homology). In this case, it is sufficient to disclose limitations.
3. Is the analytical specificity sufficiently high? If not 100%, orthogonal confirmation may need to be performed to rule out false positives OR must add limitation statement to clinical report
4. Is sufficient average coverage and minimum coverage routinely achieved?
5. Are all genes relevant to the test indication adequately covered in the majority of runs?
6. Additional questions depending on the specific test developed
G. Clinical Report Example: Methods and Limitations Statement
This assay is performed using the PerkinElmer Sciclone® G3 Workstation combined with the Agilent SureSelect Clinical Research Exome capture kit (#G9496A 5190-7344; targeting coding regions (exons) and canonical splice sites) followed by sequencing on the Illumina NextSeq 550 (High-Output v2 kit). Reads are aligned to the GRCh37 reference sequence using the Burrows-Wheeler Aligner (BWA 0.7.17), and variant calls are made using the Genomic Analysis Tool Kit (GATK v4.0.3.0). Detection of copy number variants (CNVs) encompassing 2 or more exons is performed using next-generation sequencing read data and the VisCap algorithm. CNV analysis is only performed when data meets necessary quality standards and may not be available for all cases. Variant calls are limited bioinformatically to the associated region of interest for the assay (see above for details). Sanger sequencing is used for fill in when bases have <12x coverage. All clinically significant variants are confirmed by Sanger sequencing or droplet digital PCR; variants classified as likely benign or benign are not confirmed. Variant classifications are based on ACMG/AMP criteria (Richards et al. 2015) with ClinGen rule specifications (https://www.clinicalgenome.org/working-groups/sequence-variant-interpretation/). Variants are reported according to HGVS nomenclature (www.hgvs.org/mutnomen). Likely benign and benign variants are not included in this report but are available upon request. This test does not routinely detect variants in non-coding regions (aside from the canonical splice sites), triplet repeat expansions, translocations, inversions, and copy number variants encompassing less than 2 consecutive exons. There is reduced detection for larger indels, variants in low complexity regions, and variants in regions with high homology.
This test is 99.93% sensitive (95% CI =99.92-99.94%) to detect variants changing a single base and 96.75% sensitive to detect insertion/deletions (95% CI =96.28-97.22%) within covered regions. Technical positive predictive value for single nucleotide variant changes is 99.42% (95% CI = 99.37-99.48%) and 94.16% (95% CI = 93.34-94.97%) for insertion/deletion changes within covered regions. There is demonstrated reduced detection for larger indels, especially in low complexity regions with corresponding low sequence coverage and in regions with high homology.