## Before_you_begin
Before you begin: How to use this workbook
Quality control (QC) steps are defined during test optimization and confirmed during test validation. Following test validation, organize critical QC parameters that will be monitored during each assay run once the test has been released into routine operations. Monitoring of QC data is critical to ensure that the test performs as expected, and to identify trends that may develop over time or related to certain batches of reagents. This workbook contains examples of per-run checklists as well as regularly scheduled quality assurance reports. Please note that the list is not exhaustive and will need customization for each new assay.
General: - Equipments and reagents: should be tested and qualified for quality assurance prior to clinical testing. A detailed record should be filed and kept. The lot numbers for the reagents should be documented for each testing batch. - Laboratory personnel: employee training, verification, and competency assessment should be performed before and during test performance. Personnel competency need to be verified at least semiannually during the first year of employment and at least annually thereafter. If changes occur to SOP or instrumentation, reverification should be performed.
1. Start on Tab A, Test QC Checklist
For each test, compile a list of all the items to be monitored. The parameters should be relevant to test performance, and are typically informed by the test optimization and validation runs. Some metrics will be "direct QC metrics" with defined thresholds or specifications (ie, acceptance and rejection criteria), and others will be collected in order to monitor test performance over time to be able to notice trends that require further optimization.
Preexamination QC includes items related to the specimen before the specific assay is performed (eg, specimen amount and condition). For example, if the validation studies only included whole blood EDTA specimens, and a specimen is submitted in a different type of blood tube, the laboratory should reject it as only whole blood EDTA can be processed for the ordered test according to the documented specimen rejection policy (see example Accessioning SOP (Tab C, SOP-0004)). Data for all specimens would be recorded on the "Specimen QC Log". The specimen collected not using EDTA would be further noted in the "Exception Log" and the Laboratory Supervisor notified, if there are circumstances to consider testing the specimen even though it does not meet acceptance criteria. The Supervisor would, in this case, be required to seek the Director's written approval to proceed with testing (documenting the rationale in the exception log), or the specimen would be rejected.
Examination QC includes items during performance of the specific assay. Predefined QC metrics are monitored throughout the process, ie, during library preparation, after the sequencing instrument completes the run, and during and after bioinformatic processing. Some metrics are inspected per instrument run, others for each sample. These are typically derived from the results of the test optimization and confirmed during validation, but additional trend monitoring parameters may be added over time to enhance the ability to identify the exact step in the process where an event may have caused test failure. Since germline sequencing is not quantitative, acceptance and rejection criteria are typically based on coverage-related metrics. Pre-sequencing examination metrics are helpful for identifying samples with unacceptable DNA quality or processing errors before proceeding with more costly sequencing steps. Similarly, demultiplexing metrics may identify issues with molecular barcoding or sample pooling.
For each metric, indicate:
1. The specific requirement
2. The protocol describing how to assess the specimen for the requirement
3. The relevant log to be completed by the analyst
4. Action to be taken if the specimen does not meet the requirement
5. Who can authorize processing of the specimen if the requirement is not met
2. Continue to Tab B, Test QC Monitoring
On a regular basis, the laboratory gathers all the information from the logs mentioned in the Test QC Checklist, and performs a test-specific QC review. Typically, this involves calculation of turnaround time (TAT), % positive, negative, and inconclusive reports, number of amended reports, canceled tests etc., as well as an outcomes analysis of samples noted on the exception log if specimens were authorized for processing. Other test-specific items may be added as needed to ensure accurate, timely testing and continually improve processes. In addition, a review of all test-relevant documentation is performed. This QC review is documented and signed by the Director, and relevant findings made available to all personnel.
Proficiency testing results should also be reviewed in the context of each specific test.
3. Continue to Tab C, Documentation List
For each test, compile a list of all relevant documentation. This includes all validation reports, logs, and Standard Operating Procedures (SOPs) for the test as well as any sub-assays and confirmatory tests. Indicate the minimum frequency for reviewing each document to ensure that it is up to date.
4. Glossary
Quality Control: An integral component of quality management composed of the aggregate of processes and techniques used to detect, reduce, and correct deficiencies in an analytical process. Quality control (QC) is a surveillance process in which the actions of people and performance of equipment and materials are observed in some systematic, periodic way that provides a record of consistency of performance and of action taken when performance does not conform to standards set by the laboratory. QC is a set of procedures designed to monitor the test method and the results to assure test system performance; QC includes testing control materials, charting the results and analyzing them to identify sources of error, and determining, performing, and documenting any remedial action taken as a result of this analysis.
Quality Improvement (this term has now replaced the term "Quality Assurance"): A systematic method used to identify opportunities for improvement in clinical and nonclinical systems.
Quality Management: All activities of the overall management function that determine quality policy objectives and responsibilities and the implementation of them, including the preexamination, examination, and postexamination phases of testing.
Proficiency testing: A set of tests for the continued monitoring of laboratory test performance.
© 2024 College of American Pathologists (CAP). All rights reserved.
CLSI. Human Genetic and Genomic Testing Using Traditional and High-Throughput Nucleic Acid Sequencing Methods. 3rd ed. CLSI guideline MM09. Clinical and Laboratory Standards Institute; 2023.

## A_Test_QC Checklist
Example QC Checklist
The list below is an educational example and is not meant to reflect all possible QC metrics. Please see PMID:29154853, PMID:28341590, PMID: 27993330, PMID: 24759409, PMID: 25741868, and the CAP Accreditation Checklists for additional guidance.
Section | Category | Criteria | Specific requirement. Note that these may vary between tests and laboratories | Documentation action | SOP (incl. hyperlink) | Action if requirement is not met | Metrics measured per sample and annually for continuous quality improvement (CQI) | Authorized for approval
PREEXAMINATION QC (per sample) | Specimen quality | Wrong specimen type | Whole blood | Specimen QC log (LOG-0002) | SOP-0004 (Accessioning, test A) | Inform supervisor, reject sample | YES | N/A
|  | Wrong type of tube | Purple top EDTA tube |  |  | Inform supervisor, note in exception log | YES | Director
|  | Insufficient quantity | >= 2 ml |  |  | Inform supervisor, note in exception log | YES | Supervisor
|  | Clotting (blood only) | No visible clots |  |  | Inform supervisor, note in exception log | YES | Supervisor
|  | Insufficient labeling | Labeling contains name, DOB, barcode, date of collection |  |  | Inform supervisor, note in exception log | YES | Director
|  | Expired specimen | <= 7 days since collection |  |  | Inform supervisor, note in exception log | YES | Supervisor
|  | Expired collection tube | Collection tube not expired |  |  | Inform supervisor, note in exception log | YES | Supervisor
| DNA quality and quantity | OD 260/280 ratio | >1.7 | DNA QC log (LOG-0003) | SOP-0007 (DNA QC , test A) | Inform supervisor, note in exception log | YES | Supervisor
|  | Electrophoretic analysis | Shows intact high molecular weight DNA band
|  | Quantification | >= 500 ng
EXAMINATION QC (per instrument run) | Instrument run QC | Cluster density | 865-965 k/mm2 | Test A QC log (LOG-0004) | SOP-0014 (QC acceptance criteria, test A)
|  | Base quality | Q30>=75%
| Pipeline QC | Total reads passing filter | >6 Gb
|  | % reads not assigned to any sample | <5%
| Control samples | No-template control | <500 reads aligned to target
|  | Positive control | Expected variants found
EXAMINATION QC (per sample) | Library preparation | Fragment size and distribution | >50% of fragments between 200 and 1000 bp | Test A QC log (LOG-0004) | SOP-0014 (QC acceptance criteria, test A)
|  | Library concentration | >4 nM
| Sample de-multiplexing | % reads assigned to sample | 8-14%
| Read alignment | % Reads Aligned to Target | >90%
|  | Average Target Coverage | >100X
|  | Distribution of Coverage | >95% within 20-200X
| Variant calling and annotation | Ti/Tv ratio (for exome or genome) | >2.0
| Specimen verification | Specimen identity confirmation using independent genotyping approach | >95% of tested positions on QC array match
| Data Transfer Integrity | Data transfer to cloud | FASTQ checksum matches before and after transfer to cloud storage

## B_Test_QC Improvement_Proficien
Example Quality Improvementand Proficiency Testing
Section | Category | Frequency of QA monitoring | CAP checklist requirement | Criteria | Results | Data Analysis | Implemented Corrective Action | Action plan for Continuous Quality Improvement (CQI) | SOP (incl. hyperlink)
POSTEXAMINATION QC (Specific time frame) | Reporting Statistics | Quarterly | MOL.20300 | Turn-around time (TAT), average | 13 days | Does this meet advertised TAT? Is there a positive or negative trend over the investigated period? |  | Hire additional technician for accessioning, trigger sequencing kit order when 10 kits are left (currently 7) | SOP-0005 (QA monitoring)
|  |  |  | Turn-around time (TAT), range | 3-40 days | What is the root cause for the 5% of reports with the longest TAT? What went well with the 5% of reports with the shortest TAT? Can learnings lead to improvements? |  | If possible, validate assay for lower concentrations of DNA to avoid need for concentrating samples
|  |  |  | Total number of reports | 518 | Where are the orders coming from? Can learnings lead to increase in orders? How is the trend - should capacity be increased?
|  |  |  | Positive cases (%) | 4.6% | Does this meet expectations for the test?
|  |  |  | Inconclusive cases (%) | 1.3%
|  |  |  | Negative cases (%) | 93.9%
| Amended or corrected reports |  |  | Total count of amended or corrected reports | 6 | Why did the clerical errors happen? Verify that the physician was notified. How can this be avoided? Do reports from previous batches need to be amended due to changes to variant classification/interpretation?
|  |  |  | - Due to clerical errors | 2 |  | Additional training of relevant personnel on 7/16/17. Documented on individual training records.
|  |  |  | - Due to changes to variant classification | 3
|  |  |  | - Due to changes to interpretation | 1
| Test cancellations |  |  | Total count of test cancellations | 4 | Are there any recurring patterns and trends over time? Can something be improved?
|  |  |  | - Due to wrong test ordered | 0
|  |  |  | - Due to insufficient specimen | 3
|  |  |  | - Due to degraded specimen | 1
| Batch statistics |  |  | Total number of runs | 16
|  |  |  | Number of failed runs | 1 | What caused the failure? How can a similar failure be avoided in the future? | Run failed due to PCR contamination. Implemented additional bleach cleaning before assay setup (SOP-0004)
|  |  |  | Number of runs with pipeline errors | 0
| Sample statistics |  |  | % confirmed of variants sent for confirmation | 95% | Are there recurring variants that consistently do not confirm? Should immediate changes be made to the reportable region or the pipeline? |  | Based on accumulated experience, can pipeline quality thresholds be adjusted and validated to decrease number of false positives?
|  |  |  | Number of samples with failed library preparation | 2
|  |  |  | Number of samples with failed alignment metrics | 5 | Root cause?
|  |  |  | Number of specimes/runs with abnomal GC bias | 3 | This warrant further investigation and possible re-sequencing
|  |  |  | Number of contaminated specimens | 0
| Reagent QA |  |  | Identified issues | 0
| Additional items not covered above |  |  |  |  | A sample swap was identified in batch XYZ | Additional QA step implemented and documented in SOP-0004 on 7/3/17. Further documentation in CAPA-009. | Closely monitor effectiveness of additional QC step. Re-train technicians on general quality awareness.
| Tests |  |  | Name | Date | EXAMPLE result | EXAMPLE Investigation | EXAMPLE implemented corrective action | Investigation report
PROFICIENCY TESTING | CAP proficiency test | Bi-Annually | PT: AA: COM.01500 | NGS-2017 A | Submitted 3/11/17, results received 7/21/17 | 100% on all graded responses. Educational challenges: 2 variants had incorrect HGVS notation. | Root cause for HGVS notation issue was identified in the pipeline. | Pipeline variant caller adjusted and validated. | RPT-0003
| Alternative Assessment |  |  | March 2017 Lab XYZ | submitted 2/15/17, results received 2/28/17 | 5/5 samples correct.

## C_Documentation List
Example Documentation List (TEST A)
Approval Dates
Date of test introduction | 42339
Title | Number and link | v1 | v2 | v3 | v4 | v5 | v6 | v7 | v8 | review frequency | CAP Checklist
Accessioning | SOP-0004 | v1: 2014-04-02; v2: 2014-08-10; v3: 2015-10-11; v4: 2017-02-06 |  | GEN.40825
Specimen processing | SOP-0005 | v1: 2014-09-16 |  | MOL.32390
Sample extraction | SOP-0006 | v1: 2014-09-16 |  | MOL.32525
NGS wet bench workflow, test A | SOP-0008 | v1: 2015-11-12; v2: 2017-07-03 |  | MOL.36010
Bioinformatics workflow, test A | SOP-0009 | v1: 2015-11-12 |  | MOL.36105
Sanger confirmation workflow | SOP-0013 | v1: 2014-08-06 |  | MOL.35850
Data interpretation | SOP-0015 | v1: 2014-09-10; v2: 2015-10-11 |  | MOL.36155
Specimen QC log | LOG-0002 | [9 empty] | MOL.32375
DNA QC log | LOG-0003 | [9 empty] | MOL.32430
Test A QC log | LOG-0006 | [9 empty] | MOL.34229
QC acceptance criteria, test A | SOP-0014 | [9 empty] | MOL.36147
CAPA (corrective action) | CAPA-009 | [9 empty] | GEN.20316
Validation report, test A | VAL-0017 | [9 empty] | MOL.30785
Validation report, Sanger confirmation | VAL-0003 | [9 empty] | MOL.30785
Validation report, QC Array for test A | VAL-0018 | [9 empty] | MOL.30785
Quarterly QI report Q1 2016 | RPT-0019 | [9 empty] | MOL.36020/MOL.36125
Quarterly QI report Q2 2016 | RPT-0027 | [9 empty] | MOL.36020/MOL.36125
Quarterly QI report Q3 2016 | RPT-0032 | [9 empty] | MOL.36020/MOL.36125
Quarterly QI report Q4 2016 | RPT-0037 | [9 empty] | MOL.36020/MOL.36125
Annual QI report Q1 2016-Q4 2016 | RPT-0042 | [9 empty] | MOL.36020/MOL.36125
Quarterly QI report Q1 2017 | RPT-0050 | [9 empty] | MOL.36020/MOL.36125
Quarterly QI report Q2 2017 | RTP-0056 | [9 empty] | MOL.36020/MOL.36125